4

| PREP - Safety Procedures: - Work in PCR Box or BSC - Designated CLEAN Lab coat, gloves, face mask, and face shield/goggles - Wipe down all work surfaces, tube racks, and all items used with 10% bleach solution followed by ethanol (as needed) - Turn on UV lamp for 15 minutes in PCR Box or BSC (at start or end of prev run) - Check to ensure you have sufficient amounts of all materials for number of lots to be created (Calculator) - With Scale (USSolid1), measure 28.2mL of NF Water in 50mL Falcon Tube, leave in the refrigerator > 1hr (prep ahead). - Ensure Primer Resuspension NF Water is also in refrigerator > 1 hr (prep ahead). - Do Startup of the Opentrons (last page). - Load Opentrons software, open protocol “iPGS48_x68_SOP-201-B_2.json” MAKE PGS - Thaw three tubes of each primer set for 10-30min after pulling from freezer - Spin down primer tubes after thawing - Get 50mL Falcon Tube of NF water (at least 10mL), put in rack in hood. - Add 1000uL of Primer Resuspension NF Water to each tube of primers - Do NOT stick pipette tip into tube, primers will stick to tip - Air drop (no contact) dispense into tubes, no tip change - Vortex for 10s to dissolve - Let sit 1min, then vortex again 10s - Centrifuge tubes briefly - Retrieve NF water aliquot from fridge, check that volume appears ~28.2mL - Add 600uL of 6M Guanidine-HCl to NF water in 50mL Falcon Tube - Mix by vortexing 10s - Add the full volume of all nine tubes of Primer sets to the 50mL Falcon Tube - Careful to get last bit of liquid (look) - Mix well by vortexing 30s - Remove cap and pipet any liquid in the cap and add back into the tube - Keep PGS mix in refrigerator until aliquoting or aliquot within 5min |  | PGS48 Manuf SOP Log Form Link Logged |
| --- | --- | --- |
|  | Location: |  |
|  | Name: |  |
|  | Date and Time: |  |
|  | PGS48 Lots Created: |  |
|  | Materials |  |
|  | 50mL Falcon Tube (1) - Confirm Product ID & Lot in Log Form |  |
|  | 1.5mL S.C. Tubes (67) - Confirm Product ID & Lot in Log Form |  |
|  | Nuclease-free Water (2) - Confirm Product ID & Lot in Log Form |  |
|  | FloodLAMP Primer Tubes (9) - Confirm Product ID & Lot in Log Form |  |
|  | 6M Guanidine HCl - Confirm Product ID & Lot in Log Form |  |
|  | Pipette ID - Confirm Product ID & Lot in Log Form |  |
|  | 1000uL micropipette tips - Confirm Product ID & Lot in Log Form |  |
|  | Robot tips (Opentrons 1000uL Filter Tip) |  |
|  | Equipment in AirClean Box - Mini-centrifuge - Vortex - Timer - 50mL & 1.5mL tube racks - Tip waste bin |  |

| ALIQUOT ON OPENTRONS - Arrange the robot deck according to deck layout shown below - Remove lid from tip box, verify that there is a tip in the back-left position - Fill the back and front row of each of the six tube racks with 1.5mL SC tubes, uncapped 1 - Place the 50mL PGS reservoir tube in the plate at position “8” 2 - Click “Proceed to Calibrate” button 3 - Follow on-screen instructions to complete calibration procedure 3 - Click “Start Run” - When protocol has finished, cap all tubes inside the Opentrons hood 4 - Transfer tubes to cryobox, store at -20C. - Label crybox lid with "PGS48" & Lot#. - Click “Reset Run” on OpenTrons software | 1 Make sure the caps are not obstructing the openings of the neighboring tubes. |  |  |
| --- | --- | --- | --- |
|  | 2 Ensure that the 50mL tube is fully seated in the holder with 10mL mark just showing. |  |  |
|  | 3 Calibration procedure necessary only on the first run. Subsequent runs won’t require re-calibration until the robot is disconnected / reconnected or a different program is run. Calibrate 1.5mL tube with tube out and center XY with respect to rack hole. 4 Note that the final 1-2 tubes may be short on volume. These should be either used to do QC verification of the batch or discarded. Write with a marker a "•" on the top of any short tubes. |  |  |
|  | Measure remaining volume in Falcon tube: if outside 450uL =/- 10uL, weigh all tubes indiv spec is X to Y g, reject any tubes outside spec |  |  |
| Notes: |  |  |  |

DECK LAYOUT

OpenTrons Startup

- Turn on rocker power switch in back left (-- mark).

- Press flashing blue button on front, wait until solid.

- If not already open, launch OpenTrons application with shortcut on desktop, should show .

- Start Hepa unit and put at setpoint.

- Take lid tipbox cover off.

- Ensure tip is in back left position of tipbox.

- Open json protocol.

- Run calibration.

- Leave front panel closed (down) when running.

OpenTrons Shutdown

- Place lid tipbox cover on.

- Turn off Hepa unit.

- Turn off OpenTrons rocker power switch (O mark).

- Close front panel.