| PREP - Pen & sharpie labeled with green tape - Separation of Post-Amp area - Heaters on (Heat indicator lit!) - Safety Procedures: - lab coat - gloves - face mask - face shield or goggles - Alcohol wipe sample tubes |  | Log Link: SOP-104-B_1 {SITE} |  |
| --- | --- | --- | --- |
|  |  | Location: |  |
|  | Name: |  |  |
|  | Date: Time: |  |  |

CONTAMINATION INFRACTIONS

- PPE lapse/removal

- Use green labeled pen/marker with bare hands and not putting it to be cleaned afterwards

- Use unmarked pen/marker with clean gloves (without cleaning gloves after or changing gloves)

- Bare hands on anything improper, except in desk area

- Reuse pipette tip improperly

- Touch pipette tip to anything and then not discard it

- Touch anything dirty then go back to work (face, mask, uncleaned phone, uncleaned laptop)

# NOTE

It is recommended that 1XISS and Reaction Mix be prepared before the assay run (instructions below).

DOCUMENTATION

- Check off each item after completing it

- Write legibly

# GLOSSARY OF TERMS

100XIS…………………………..	100-fold concentrated Inactivation Solution

1XISS…………………………….	1X Inactivation Saline Solution

PGS……………………………….	Primer Guanidine Solution

CLAMP MM………………….	Colorimetric Loop-Mediated Isothermal Amplification Master Mix

RM…………………………………	Reaction Mix (CLAMP MM and PGS)

TPC……………………………….	Twist Positive Control

| 1XISS PREP - Record volumes of saline & 100XIS in 1XISS Prep form prior to pipetting - Combine saline and 100XIS in a 50mL falcon tube or 30mL Chub tube, depending on volume - Vortex 30 seconds - Label tube with 1XISS ID (e.g., “1X_MMDD-1”) and record the same ID into the 1XISS Prep Form and the table on this form | Sample Batch ID (from App): |  |  |
| --- | --- | --- | --- |
|  | Inactiv Heater: |  |  |
|  | 1X Inactiv Saline Soln ID: |  |  |
|  | Dispense: Pipette or Manual Disp or Electric Disp |  |  |

INACTIVATION

- Add 1mL of 1XISS to each sample tube

- Verify that tubes are properly sealed afterward (i.e., not cross-threaded)

- Vortex 10 sec (if individual tubes) or 30 sec (if tubes-in-rack)

- Heat 8 min in Water Bath (set to 99°C)

- Cool 10 min

- Put in refrigerator before 30 minute mark

REACTION MIX PREP

- Complete the entry in the RM Prep Form prior to pipetting

- Calc volumes correctly for PGS and CLAMP based on number of reactions

- Remove CLAMP MM and PGS from freezer to thaw, but do not allow them to warm

- Briefly spin down reagent tubes (removes reagent from the cap and tube walls)

- Pipette CLAMP MM and PGS into a 1.5mL tube using calculated volumes

- Promptly return PGS and CLAMP source tubes to freezer

- Vortex and spin down the RM tube

- Aliquot 23uL of RM into each reaction tube (plate or strips-of-8)

  - Pipetting at the bottom of tubes

  - No blowout while dispensing reaction mix (to prevent bubbles in bottom of well)

- Visually check the volumes in each reaction tube (that they appear to be even)

- If using a plate, cover filled tubes with rack lid to prevent dust

- At this stage, RM reaction tubes can be capped and placed in the freezer for storage if they are not to be used fresh.

| AMPLIFICATION REACTION - Remove PCR cold block from freezer ~10 min before adding tubes - Prep RM plate/strips (thawed, not warm) - Set up tips to align with reactions - Use a lookup rack to keep track of sample order - Tip layout aligns with arrangement of sample tubes in lookup rack & aligns with layout of reaction plate/strips - Add negative control first (2uL 1XISS) - Add 2uL of each inactivated sample to reaction tubes, visually checking the volume in the tip each time. Pipette up & down 5X, blow out in liquid, tip touch. | Amp Heater: |  | Strip8 Tubes or Plates |
| --- | --- | --- | --- |
|  | Reaction Mix: Frozen or Fresh |  |  |
|  | Reaction Mix ID: |  |  |
|  | Num Reactions (including controls): |  |  |
|  | Pos Ctrl ID: |  |  |
|  | Neg Ctrl ID (today's 1XISS): |  |  |
|  | Initial Inconclusives?: |  |  |
| - With each add, verify the correct position (tip vs sample tube vs reaction tube) - Transfer plate/strip to PCR rack (not cold block - rxns don’t go to amp cold) - Retrieve TPC from freezer, touching only with one hand (consider this hand dirty) - Add 2uL TPC to pos. ctrl, return TPC to freezer (IMMEDIATELY change glove) - Cap reaction tubes / seal plate - Change gloves again after loading tubes/plate onto amp heater - Set main timer for 25 min - Set secondary timer for 24 min - Put sample tubes in fridge - Remove rxns from amp promptly, wait 1 minute before imaging - Take photo in lightbox - For photo apply crop and ‘Vivid’ filter (or increase saturation 70% for android) - Log amp run sheet with QR code - On app: Intake, Process and Result tubes | Notes: |  |  |

|  | 1 | 2 | 3 | 4 | 5 | 6 |
| --- | --- | --- | --- | --- | --- | --- |
| A 1 |  |  |  |  |  |  |
| B 2 |  |  |  |  |  |  |
| C 3 |  |  |  |  |  |  |
| D 4 |  |  |  |  |  |  |
| E 5 |  |  |  |  |  |  |
| F 6 |  |  |  |  |  |  |
| G 7 |  |  |  |  |  |  |
| H 8 |  |  |  |  |  |  |

|  | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- |
| A 1 |  |  |  |  |  |  |
| B 2 |  |  |  |  |  |  |
| C 3 |  |  |  |  |  |  |
| D 4 |  |  |  |  |  |  |
| E 5 |  |  |  |  |  |  |
| F 6 |  |  |  |  |  |  |
| G 7 |  |  |  |  |  |  |
| H 8 |  |  |  |  |  |  |