| SOP - Only approved personnel can perform this assessment - Watch video and check items - Note timestamp of any deviations - Determine PASS or FAIL - Log assessment with Form |  | Certification Run Assessment Form Link | version 1.1 |
| --- | --- | --- | --- |
|  |  | Video Link: |  |
|  | Assessor Name: |  |  |
|  | Date and Time: |  |  |
|  | PASS or FAIL |  |  |

Safety

- Proper PPE (Gloves, Lab Coat, N-95 Mask, and Goggles / Face Shield)

Contamination

- Pen and sharpie labeled with green tape

- Separation of Post-Amp area

Contamination Infractions

- PPE lapse/removal

- Use green labeled pen/marker with bare hands and not putting it to be cleaned afterwards

- Use unmarked pen/marker with clean gloves (without cleaning gloves after or changing gloves)

- Bare hands on anything improper, except in desk area

- Reuse pipette tip improperly

- Touch pipette tip to anything and then not discard it

- Touch anything dirty then go back to work (face, mask, uncleaned phone, uncleaned laptop)

Prep

- Turn on heaters, heat indicating light

- Alcohol wipe the collection tubes

Documentation

- Read each item then check it off after completing it

- Writing legible

Inactivation

- Write down on log sheet the amounts of saline and 100X prior to pipetting

- Label 1XISS and logging the preparation correctly

- Volumes correct for Saline and 100X IS

- Mixing correct for 1XISS

- Volume correct for 1XISS added to samples (1mL)

- Sample tubes opened and closed properly

- Vortex of samples correct

- Timing correct for inactivation heating and cooling

Reaction Mix Prep (strip8 tubes)

- Write down entry on log sheet prior to starting prep

- Calc volumes correctly for PGS and CLAMP based on number of reactions

- CLAMP and PGS thawed properly, so not to get too warm

- Spin down reagents and reaction mix

- Correct volumes for CLAMP and PGS

- Get PGS and CLAMP source tubes back in freezer quickly

- Vortex and spin down reaction mix tube

- Marking the top tube in strip with 1,2,3, etc

- Correct volume of reaction mix in each pcr tube (23uL)

- Pipetting at the bottom of tubes

- No blowout while dispensing reaction mix

- Volume check strip tubes

- Keep reactions covered strip tubes with rack lid to prevent dust

Amp Run (strip8 tubes)

- Get Cold PCR Block out ~10min before adding tubes

- Reaction mix tubes ready (thawed but not too warm (<10 min)

- Strip 8 tubes labeled on top tube of each strip

- Tips setup properly to align with strips8 tubes to be used

- Add negative control 1st

- Using a lookup rack to keep continuity of sample order

- Visually checking the 2uL in the tip from each inactivated sample

- Amp pipetting technique correct - pipet up and down 5 times, blowout in liquid, tip touch

- Well position correct and correspondence between tip, strip tube# and, where sample tube is put in lookup rack

- Transfer strip to regular PCR block (so they aren’t cold going to Amp)

- Only touching positive control (TPC) with one hand

- Glove change after putting the positive control (TPC) back in freezer and before caps

- Not touching the top of strip8 tubes or inside of a strip8 cap

- Completely sealed strip tubes - check at eye level

- Change gloves again after touching loading tubes/plate onto amp heater

- Set main timer for 25min

- Set phone timer for 24min

- Put sample tubes in fridge during amp heating

- Wait at least 1 minute after pulling off the amp heater

- Take photo in lightbox

- For photo apply crop and vivid filter

- Logging amp run sheet properly with google form entry

App

- Intake tubes correctly

- Process and Result tubes correctly

ONLY COMPLETE THIS LAST SECTION IF TRAINING FOR PLATE SCALE TESTING

Dispenser Demonstration

- Unpack an electric dispenser

- Properly prime the dispenser

- Dispense 1mL into tube without touching the tip

Amp Run (Abbreviated Plate)

- Prepare reaction Mix following all previous guidelines

- Pipette Reaction Mix into a reservoir

- Use Multichannel Pipette to fill Column 1

- Fully seal plate with foil seal

- Pierce B1 to add negative control

- Pierce C1 - H1 to add samples

- Pierce A1 to add TPC, following all TPC guidelines

- Glove change

- Use a second foil seal on top of the first

- Tight seal, no delamination

- Put on Amp heater

- Glove change