CDC 2019-Novel Coronavirus (2019-nCoV)

Real-Time RT-PCR Diagnostic Panel

For Emergency Use Only

Instructions for Use

Catalog # 2019-nCoVEUA-01

1000 reactions

For In-vitro Diagnostic (IVD) Use

Rx Only

Centers for Disease Control and Prevention

Division of Viral Diseases

1600 Clifton Rd NE

Atlanta GA 30329

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table of Contents

Intended Use ............................................................................................................................. 3  Summary and Explanation....................................................................................................... 4  Principles of the Procedure ..................................................................................................... 4  Materials Required (Provided)................................................................................................. 6  Materials Required (But Not Provided) ................................................................................... 7  Warnings and Precautions .................................................................................................... 10  Reagent Storage, Handling, and Stability............................................................................. 12  Specimen Collection, Handling, and Storage....................................................................... 12  Specimen Referral to CDC ..................................................................................................... 13  Reagent and Controls Preparation........................................................................................ 13  General Preparation ............................................................................................................... 14  Nucleic Acid Extraction.......................................................................................................... 14  Assay Set Up........................................................................................................................... 16

Create a Run Template on the Applied Biosystems 7500 Fast Dx Real-time PCR  Instrument (Required if no template exists)......................................................................... 21

Defining the Instrument Settings .......................................................................................... 27  Running a Test........................................................................................................................ 29  Interpretation of Results and Reporting ............................................................................... 35  2019-nCoV rRT-PCR Diagnostic Panel Results Interpretation Guide ................................ 38  Quality Control........................................................................................................................ 39  Limitations .............................................................................................................................. 39  Conditions of Authorization for the Laboratory................................................................... 40  Performance Characteristics................................................................................................. 42  Disposal................................................................................................................................... 52  References .............................................................................................................................. 52  Revision History ..................................................................................................................... 53  Contact Information, Ordering, and Product Support ......................................................... 53

Appendix A: Heat Treatment Alternative to Extraction ....................................................... 54

1

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Appendix B: Pooled Specimen Preparation and Processing ............................................. 58  Appendix C: Implementation and Monitoring of Pooled Specimen Testing...................... 67

2

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Intended Use

The CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel is a real-time RT-PCR  test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in upper and lower  respiratory specimens (such as nasopharyngeal or oropharyngeal swabs, sputum, lower respiratory tract  aspirates, bronchoalveolar lavage, and nasopharyngeal wash/aspirate or nasal aspirate) collected from  individuals suspected of COVID-19 by their healthcare provider1.

This test is also for the qualitative detection of nucleic acid from the SARS-CoV-2 in pooled samples

containing up to four of the individual upper respiratory swab specimens (nasopharyngeal (NP),  oropharyngeal (OP), NP/OP combined, or nasal swabs) that were collected using individual vials  containing transport media from individuals suspected of COVID-19 by their healthcare provider.  Negative results from pooled testing should not be treated as definitive. If a patient’s clinical signs and  symptoms are inconsistent with a negative result or results are necessary for patient management, then

the patient should be considered for individual testing. Specimens included in pools with a positive,

inconclusive, or invalid result must be tested individually prior to reporting a result. Specimens with low

viral loads may not be detected in sample pools due to the decreased sensitivity of pooled testing.

Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of

1988 (CLIA), 42 U.S.C. § 263a, that meet the requirements to perform high complexity tests.

Results are for the identification of SARS-CoV-2 RNA. SARS-CoV-2 RNA is generally detectable in upper

and lower respiratory specimens during infection. Positive results are indicative of active infection with  SARS-CoV-2 but do not rule out bacterial infection or co-infection with other viruses. The agent detected  may not be the definite cause of disease. Laboratories within the United States and its territories are

required to report all results to the appropriate public health authorities.

Negative results do not preclude SARS-CoV-2 infection and should not be used as the sole basis for  treatment or other patient management decisions. Negative results must be combined with clinical  observations, patient history, and epidemiological information.

Testing with the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel is intended for use by trained  laboratory personnel who are proficient in performing real-time RT-PCR assays. The CDC 2019-Novel  Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel is only for use under a Food and Drug  Administration’s Emergency Use Authorization.

1 For this EUA, a healthcare provider includes, but is not limited to, physicians, nurses, pharmacists, technologists, laboratory  directors, epidemiologists, or any other practitioners or allied health professionals.

3

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Summary and Explanation

An outbreak of pneumonia of unknown etiology in Wuhan City, Hubei Province, China was initially  reported to WHO on December 31, 2019. Chinese authorities identified a novel coronavirus (2019-nCoV,  also referred to as SARS-CoV-2), which has resulted in millions of confirmed human infections globally.  Cases of asymptomatic infection, mild illness, severe illness, and deaths have been reported.

The CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel is a molecular in vitro diagnostic test that aids in  the detection and diagnosis of SARS-CoV-2 infection and is based on widely used nucleic acid  amplification technology. The product contains oligonucleotide primers and dual-labeled hydrolysis

probes (TaqMan®) and control material used in rRT-PCR for the in vitro qualitative detection of 2019- nCoV RNA in respiratory specimens.

The term “qualified laboratories” refers to laboratories in which all users, analysts, and any person

reporting results from use of this device should be trained to perform and interpret the results from this

procedure by a competent instructor prior to use.

Principles of the Procedure

The oligonucleotide primers and probes for detection of 2019-nCoV were selected from regions of the

virus nucleocapsid (N) gene. The panel is designed for specific detection of SARS-CoV-2 (two

primer/probe sets). An additional primer/probe set to detect the human RNase P gene (RP) in control  samples and clinical specimens is also included in the panel.

RNA isolated and purified from upper and lower respiratory specimens is reverse transcribed to cDNA

and subsequently amplified in the Applied Biosystems 7500 Fast Dx Real-Time PCR Instrument with SDS

version 1.4 software. In the process, the probe anneals to a specific target sequence located between

the forward and reverse primers. During the extension phase of the PCR cycle, the 5’ nuclease activity of

Taq polymerase degrades the probe, causing the reporter dye to separate from the quencher dye,

generating a fluorescent signal. With each cycle, additional reporter dye molecules are cleaved from

their respective probes, increasing the fluorescence intensity. Fluorescence intensity is monitored at  each PCR cycle by Applied Biosystems 7500 Fast Dx Real-Time PCR System with SDS version 1.4 software.

Detection of viral RNA not only aids in the diagnosis of illness but also provides epidemiological and

surveillance information.

4

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Summary of Preparation and Testing Process

Upon receipt of rRT-PCR Panel  reagents

Upon obtaining

Resuspend

primer/probe mix,  aliquot and store at  ≤ -20°C

Resuspend and  aliquot nCoVPC,  store at -70°C

sample Extract sample RNA

and HSC RNA

Prepare master mix

(15 ∝L)

Prepare rRT-PCR

plate (5 ∝L RNA)

Run assay on

ABI 7500Fast Dx

Analyze data

Report results

5

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Materials Required (Provided)

Note: CDC will maintain on its website a list of commercially available lots of primer and probe sets

and/or positive control materials that are acceptable alternatives to the CDC primer and probe set

and/or positive control included in the Diagnostic Panel. Only material distributed through the CDC  International Reagent Resource and specific lots of material posted to the CDC website are acceptable  for use with this assay under CDC’s Emergency Use Authorization.

This list of acceptable alternative lots of primer and probe materials and/or positive control materials

will be available at:

https://www.cdc.gov/coronavirus/2019-nCoV/lab/virus-requests.html

Primers and Probes:

Catalog #2019-nCoVEUA-01 Diagnostic Panel Box #1:

| Reagent Label | Part # | Description | Quantity / Tube | Reactions / Tube |
| --- | --- | --- | --- | --- |
| 2019-nCoV_N1 | RV202001 RV202015 | 2019-nCoV_N1 Combined Primer/Probe Mix | 22.5 nmol | 1000 |
| 2019-nCoV_N2 | RV202002 RV202016 | 2019-nCoV_N2 Combined Primer/Probe Mix | 22.5 nmol | 1000 |
| RP | RV202004 RV202018 | Human RNase P Combined Primer/Probe Mix | 22.5 nmol | 1000 |

Positive Control (either of the following products are acceptable):

Catalog #2019-nCoVEUA-01 Diagnostic Panel Box #2:

| Reagent Label | Part # | Description | Quantity | Notes |
| --- | --- | --- | --- | --- |
| nCoVPC | RV202005 | 2019-nCoV Positive Control (nCoVPC) For use as a positive control with the CDC 2019- nCoV Real-Time RT-PCR Diagnostic Panel procedure. The nCoVPC contains noninfectious positive control material supplied in a dried state and must be resuspended before use. nCoVPC consists of in vitro transcribed RNA. nCoVPC will yield a positive result with each assay in the 2019-nCoV Real-Time RT-PCR Diagnostic Panel including RP. | 4 tubes | Provides (800) 5 µL test reactions |

6

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Catalog #VTC-04 CDC 2019-nCoV Positive Control (nCoVPC)

| Reagent Label | Part # | Description | Quantity | Notes |
| --- | --- | --- | --- | --- |
| nCoVPC | RV202005 | 2019-nCoV Positive Control (nCoVPC) For use as a positive control with the CDC 2019- nCoV Real-Time RT-PCR Diagnostic Panel procedure. The nCoVPC contains noninfectious positive control material supplied in a dried state and must be resuspended before use. nCoVPC consists of in vitro transcribed RNA. nCoVPC will yield a positive result with each assay in the 2019-nCoV Real-Time RT-PCR Diagnostic Panel including RP. | 4 tubes | Provides (800) 5 µL test reactions |

Materials Required (But Not Provided)

Human Specimen Control (HSC)

| Description | Quantity | CDC Catalog No. |
| --- | --- | --- |
| Manufactured by CDC. For use as a nucleic acid extraction procedural control to demonstrate successful recovery of nucleic acid as well as extraction reagent integrity. The HSC consists of noninfectious (beta Propiolactone treated) cultured human cell material supplied as a liquid suspended in 0.01 M PBS at pH 7.2-7.4. | 10 vials x 500 µL | KT0189 |

Acceptable alternatives to HSC:

• Negative human specimen material: Laboratories may prepare a volume of human specimen

material (e.g., human sera or pooled leftover negative respiratory specimens) to extract and

run alongside clinical samples as an extraction control. This material should be prepared in  sufficient volume to be used across multiple runs. Material should be tested prior to use as

the extraction control to ensure it generates the expected results for the HSC listed in these  instructions for use.

• Contrived human specimen material: Laboratories may prepare contrived human specimen  materials by suspending any human cell line (e.g., A549, Hela, or 293) in PBS. This material

should be prepared in sufficient volume to be used across multiple runs. Material should be

tested prior to use as the extraction control to ensure it generates the expected results for  the HSC listed in these instructions for use.

CDC will maintain on its website a list of commercially alternative extraction controls, if applicable,  that are acceptable for use with this assay under CDC’s Emergency Use Authorization, at:  https://www.cdc.gov/coronavirus/2019-nCoV/lab/virus-requests.html

7

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

rRT-PCR Enzyme Mastermix Options

| Reagent | Quantity | Catalog No. |
| --- | --- | --- |
| Quantabio qScript XLT One-Step RT-qPCR ToughMix | 100 x 20 μL rxns (1 x 1 mL) | 95132-100 |
|  | 2000 x 20 μL rxns (1 x 20 mL) | 95132-02K |
|  | 500 x 20 μL rxns (5 x 1 mL) | 95132-500 |
| Quantabio UltraPlex 1-Step ToughMix (4X) | 100 x 20 µL rxns (500 µL) | 95166-100 |
|  | 500 x 20 μL rxns (5 x 500 µL) | 95166-500 |
|  | 1000 x 20 μL rxns (1 x 5 mL) | 95166-01K |
| Promega GoTaq® Probe 1- Step RT-qPCR System | 200 x 20 μL rxns (2 mL) | A6120 |
|  | 1250 x 20 μL rxns (12.5 mL) | A6121 |
| Thermofisher TaqPath™ 1-Step RT-qPCR Master Mix, CG | 1000 reactions | A15299 |
|  | 2000 reactions | A15300 |

RNA Extraction Options

For each of the kits listed below, CDC has confirmed that the external lysis buffer is effective for  inactivation of SARS-CoV-2.

| Instrument/Manufacturer | Extraction Kit | Catalog No. |
| --- | --- | --- |
| QIAGEN | 2QIAmp DSP Viral RNA Mini Kit | 50 extractions (61904) |
|  | 2QIAamp Viral RNA Mini Kit | 50 extractions (52904) 250 extractions (52906) |
| QIAGEN EZ1 Advanced XL | 2EZ1 DSP Virus Kit | 48 extractions (62724) Buffer AVL (19073 or 19089) EZ1 Advanced XL DSP Virus Card (9018703) |
|  | 2EZ1 Virus Mini Kit v2.0 | 48 extractions (955134) Buffer AVL (19073 or 19089) EZ1 Advanced XL Virus Card v2.0 (9018708) |
| Roche MagNA Pure 24 | 2MagNA Pure 24 Total NA Isolation Kit | 96 extractions (07 658 036 001) External Lysis Buffer (06 374 913 001, 12 239 469 103, 03 246 779 001 or 03 246 752 001) |
| Roche MagNA Pure 96 | 2DNA and Viral NA Small Volume Kit | 576 extractions (06 543 588 001) External Lysis Buffer (06 374 913 001, 12 239 469 103, 03 246 779 001 or 03 246 752 001) |
| 1Roche MagNA Pure LC | 2Total Nucleic Acid Kit | 192 extractions (03 038 505 001) |

8

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

| Instrument/Manufacturer | Extraction Kit | Catalog No. |
| --- | --- | --- |
| 1Roche MagNA Pure Compact | 2Nucleic Acid Isolation Kit I | 32 extractions (03 730 964 001) |
| Promega Maxwell® RSC 48 and Maxwell® CSC 48 | 3Maxwell® RSC Viral Total Nucleic Acid Purification Kit | 48 extractions (AS1330) 144 extractions (ASB1330) |
| 1QIAGEN QIAcube | 2QIAmp DSP Viral RNA Mini Kit | 50 extractions (61904) |
|  | 2QIAamp Viral RNA Mini Kit | 50 extractions (52904) 250 extractions (52906) |
| 1, 3bioMérieux NucliSENS® easyMAG® and 1, 3bioMérieux EMAG® (Automated magnetic extraction reagents sold separately. Both instruments use the same reagents and disposables, with the exception of tips.) |  | EasyMAG® Magnetic Silica (280133) EasyMAG® Lysis Buffer (280134) EasyMAG® Lysis Buffer, 2 mL (200292) EasyMAG® Wash Buffers 1,2, and 3 (280130, 280131, 280132) EasyMAG® Disposables (280135) Biohit Pipette Tips (easyMAG® only) (280146) EMAG®1000μL Tips (418922) |

1Equivalence and performance of these extraction platforms for extraction of viral RNA were demonstrated with the CDC  Human Influenza Virus Real-Time RT-PCR Diagnostic Panel (K190302). Performance characteristics of these extraction  platforms with 2019-nCoV (SARS CoV-2) have not been demonstrated.

2 CDC has confirmed that the external lysis buffer used with this extraction method is effective for inactivation of SARS CoV-2.

3 CDC has compared the concentration of inactivating agent in the lysis buffer used with this extraction method and has  determined the concentration to be within the range of concentrations found effective in inactivation of SARS-CoV-2.

Alternative to Extraction:

If a laboratory cannot access adequate extraction reagents to support testing demand due to the  global shortage of reagents, CDC has evaluated a heat treatment procedure for upper respiratory

specimens using the Quantabio UltraPlex 1-Step ToughMix (4X), CG. Though performance was  comparable, this method has been evaluated with a limited number of clinical specimens and a  potential reduction in sensitivity due to carryover of inhibitory substances or RNA degradation cannot  be ruled out. It should only be used when a jurisdiction determines that the testing need is great  enough to justify the risk of a potential loss of sensitivity. Heat-treated specimens generating  inconclusive or invalid results should be extracted with an authorized extraction method prior to

retesting. Details and procedure for the heat treatment alternative to extraction may be found in

Appendix A.

Equipment and Consumables Required (But Not Provided)

▪ Vortex mixer

▪ Microcentrifuge

▪ Micropipettes (2 or 10 μL, 200 μL and 1000 μL)

▪ Multichannel micropipettes (5-50 μL)

9

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

▪ Racks for 1.5 mL microcentrifuge tubes

▪ 2 x 96-well -20°C cold blocks

▪ 7500 Fast Dx Real-Time PCR Systems with SDS 1.4 software (Applied Biosystems; catalog

#4406985 or #4406984)

▪ Extraction systems (instruments): QIAGEN EZ1 Advanced XL, QIAGEN QIAcube, Roche MagNA  Pure 24, Roche MagNA Pure 96, Promega Maxwell® RSC 48, Roche MagNA Pure LC, Roche

MagNA Pure Compact, bioMérieux easyMAG, and bioMérieux EMAG

▪ Molecular grade water, nuclease-free

▪ 10% bleach (1:10 dilution of commercial 5.25-6.0% hypochlorite bleach)  ▪ DNAZapTM (Ambion, cat. #AM9890) or equivalent

▪ RNase AWAY™ (Fisher Scientific; cat. #21-236-21) or equivalent

▪ Disposable powder-free gloves and surgical gowns

▪ Aerosol barrier pipette tips

▪ 1.5 mL microcentrifuge tubes (DNase/RNase free)

▪ 0.2 mL PCR reaction plates (Applied Biosystems; catalog #4346906 or #4366932)  ▪ MicroAmp Optical 8-cap Strips (Applied Biosystems; catalog #4323032)

Qualifying Alternative Components:

If a laboratory modifies this test by using unauthorized, alternative components (e.g., extraction

methods or PCR instruments), the modified test is not authorized under this EUA. FDA’s Policy for  Diagnostic Tests for Coronavirus Disease-2019 during the Public Health Emergency, updated May 11,

2020, does not change this. As part of this policy, FDA does not intend to object when a laboratory  modifies an EUA-authorized test, which could include using unauthorized components, without

obtaining an EUA or EUA amendment, where the modified test is validated using a bridging study to the

EUA-authorized test.

Warnings and Precautions

• For in vitro diagnostic use (IVD).

• This test has not been FDA cleared or approved; this test has been authorized by FDA under

an EUA for use by laboratories certified under CLIA, 42 U.S.C. § 263a, that meet

requirements to perform high complexity tests.

• This test has been authorized only for the detection of nucleic acid from SARS CoV-2, not  for any other viruses or pathogens.

• This test is only authorized for the duration of the declaration that circumstances exist  justifying the authorization of emergency use of in vitro diagnostic tests for detection  and/or diagnosis of COVID-19 under Section 564(b)(1) of the Federal Food, Drug and

Cosmetic Act, 21 U.S.C. § 360bbb-3(b)(1), unless the authorization is terminated or revoked  sooner.

• Follow standard precautions. All patient specimens and positive controls should be considered  potentially infectious and handled accordingly.

• Do not eat, drink, smoke, apply cosmetics or handle contact lenses in areas where reagents and  human specimens are handled.

10

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

• Handle all specimens as if infectious using safe laboratory procedures. Refer to Interim

Laboratory Biosafety Guidelines for Handling and Processing Specimens Associated with 2019- nCoV https://www.cdc.gov/coronavirus/2019-nCoV/lab-biosafety-guidelines.html.  • Specimen processing should be performed in accordance with national biological safety

regulations.

• If infection with 2019-nCoV is suspected based on current clinical and epidemiological screening  criteria recommended by public health authorities, specimens should be collected with  appropriate infection control precautions.

• Performance characteristics have been determined with human upper respiratory specimens and

lower respiratory tract specimens from human patients with signs and symptoms of respiratory  infection.

• Perform all manipulations of live virus samples within a Class II (or higher) biological safety  cabinet (BSC).

• Use personal protective equipment such as (but not limited to) gloves, eye protection, and lab

coats when handling kit reagents while performing this assay and handling materials including  samples, reagents, pipettes, and other equipment and reagents.

• Amplification technologies such as PCR are sensitive to accidental introduction of PCR product  from previous amplifications reactions. Incorrect results could occur if either the clinical  specimen or the real-time reagents used in the amplification step become contaminated by  accidental introduction of amplification product (amplicon). Workflow in the laboratory should  proceed in a unidirectional manner.

▪ Maintain separate areas for assay setup and handling of nucleic acids.

▪ Always check the expiration date prior to use. Do not use expired reagents. Do not

substitute or mix reagents from different kit lots or from other manufacturers.

▪ Change aerosol barrier pipette tips between all manual liquid transfers.

▪ During preparation of samples, compliance with good laboratory techniques is essential

to minimize the risk of cross-contamination between samples and the inadvertent  introduction of nucleases into samples during and after the extraction procedure. Proper

aseptic technique should always be used when working with nucleic acids.

▪ Maintain separate, dedicated equipment (e.g., pipettes, microcentrifuges) and supplies  (e.g., microcentrifuge tubes, pipette tips) for assay setup and handling of extracted  nucleic acids.

▪ Wear a clean lab coat and powder-free disposable gloves (not previously worn) when

setting up assays.

▪ Change gloves between samples and whenever contamination is suspected.  ▪ Keep reagent and reaction tubes capped or covered as much as possible.

▪ Primers, probes (including aliquots), and enzyme master mix must be thawed and

maintained on a cold block at all times during preparation and use.

▪ Work surfaces, pipettes, and centrifuges should be cleaned and decontaminated with  cleaning products such as 10% bleach, DNAZap™, or RNase AWAY™ to minimize risk of

nucleic acid contamination. Residual bleach should be removed using 70% ethanol.  • RNA should be maintained on a cold block or on ice during preparation and use to ensure  stability.

• Dispose of unused kit reagents and human specimens according to local, state, and federal

regulations.

11

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Reagent Storage, Handling, and Stability

• Store all dried primers and probes and the positive control, nCoVPC, at 2-8°C until re-hydrated for

use. Store liquid HSC control materials at ≤ -20°C.

Note: Storage information is for CDC primer and probe materials obtained through the

International Reagent Resource. If using commercial primers and probes, please refer to the

manufacturer’s instructions for storage and handling.

• Always check the expiration date prior to use. Do not use expired reagents.

• Protect fluorogenic probes from light.

• Primers, probes (including aliquots), and enzyme master mix must be thawed and kept on a cold  block at all times during preparation and use.

• Do not refreeze probes.

• Controls and aliquots of controls must be thawed and kept on ice at all times during preparation

and use.

Specimen Collection, Handling, and Storage

Inadequate or inappropriate specimen collection, storage, and transport are likely to yield false test

results. Training in specimen collection is highly recommended due to the importance of specimen

quality. CLSI MM13-A may be referenced as an appropriate resource.

⮚ Collecting the Specimen

• Refer to Interim Guidelines for Collecting, Handling, and Testing Clinical Specimens for COVID 19 https://www.cdc.gov/coronavirus/2019-nCoV/guidelines-clinical-specimens.html

• Follow specimen collection device manufacturer instructions for proper collection methods.  • Swab specimens should be collected using only swabs with a synthetic tip, such as nylon or  Dacron®, and an aluminum or plastic shaft. Calcium alginate swabs are unacceptable and  cotton swabs with wooden shafts are not recommended. Place swabs immediately into  sterile tubes containing 1-3 mL of appropriate transport media, such as viral transport media

(VTM).

⮚ Transporting Specimens

• Specimens must be packaged, shipped, and transported according to the current edition of  the International Air Transport Association (IATA) Dangerous Goods Regulation. Follow  shipping regulations for UN 3373 Biological Substance, Category B when sending potential  2019-nCoV specimens. Store specimens at 2-8°C and ship overnight to CDC on ice pack. If a  specimen is frozen at -70°C or lower, ship overnight to CDC on dry ice.

⮚ Storing Specimens

• Specimens can be stored at 2-8°C for up to 72 hours after collection.

• If a delay in extraction is expected, store specimens at -70°C or lower.

• Extracted nucleic acid should be stored at -70°C or lower.

12

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Specimen Referral to CDC

For state and local public health laboratories:

• Ship all specimens overnight to CDC.

• Ship frozen specimens on dry ice and non-frozen specimens on cold packs.

• Refer to the International Air Transport Association (IATA - www.iata.org) for requirements  for shipment of human or potentially infectious biological specimens. Follow shipping  regulations for UN 3373 Biological Substance, Category B when sending potential 2019-nCoV

specimens.

• Prior to shipping, notify CDC Division of Viral Diseases (see contact information below) that

you are sending specimens.

• Send all samples to the following recipient:

Centers for Disease Control and Prevention

c/o STATT

Attention: Unit 66

1600 Clifton Rd., Atlanta, GA 30329-4027

Phone: (404) 639-3931

The emergency contact number for CDC Emergency Operations Center (EOC) is

770-488-7100.

All other laboratories that are CLIA certified and meet requirements to perform high complexity  testing:

• Please notify your state and/or local public health laboratory for specimen referral and  confirmatory testing guidance.

Reagent and Controls Preparation

NOTE: Storage information is for materials obtained through the CDC International Reagent Resource.

If using commercial products for testing, please refer to the manufacturer’s instructions for storage,  handling, and preparation instructions.

Primer and Probe Preparation:

1) Upon receipt, store dried primers and probes at 2-8°C.

2) Precautions: These reagents should only be handled in a clean area and stored at appropriate  temperatures (see below) in the dark. Freeze-thaw cycles should be avoided. Maintain cold  when thawed.

3) Using aseptic technique, suspend dried reagents in 1.5 mL of nuclease-free water and allow

to rehydrate for 15 min at room temperature in the dark.

4) Mix gently and aliquot primers/probe in 300 μL volumes into 5 pre-labeled tubes. Store a

single, working aliquot of primers/probes at 2-8°C in the dark. Store remaining aliquots at  ≤ -20°C in a non-frost-free freezer. Do not refreeze thawed aliquots (stable for up to 4 months  at 2-8°C).

13

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

2019-nCoV Positive Control (nCoVPC) Preparation:

1) Precautions: This reagent should be handled with caution in a dedicated nucleic acid handling  area to prevent possible contamination. Freeze-thaw cycles should be avoided. Maintain on  ice when thawed.

2) Resuspend dried reagent in each tube in 1 mL of nuclease-free water to achieve the proper

concentration. Make single use aliquots (approximately 30 μL) and store at ≤ -70°C.

3) Thaw a single aliquot of diluted positive control for each experiment and hold on ice until  adding to plate. Discard any unused portion of the aliquot.

Human Specimen Control (HSC) (not provided):

1) Human Specimen Control (HSC) or one of the listed acceptable alternative extraction controls  must be extracted and processed with each specimen extraction run.

2) Refer to the Human Specimen Control (HSC) package insert for instructions for use.  No Template Control (NTC) (not provided):

1) Sterile, nuclease-free water

2) Aliquot in small volumes

3) Used to check for contamination during specimen extraction and/or plate set-up  General Preparation

Equipment Preparation

Clean and decontaminate all work surfaces, pipettes, centrifuges, and other equipment prior to use.

Decontamination agents should be used including 10% bleach, 70% ethanol, and DNAzap™, or RNase  AWAY™ to minimize the risk of nucleic acid contamination.

Nucleic Acid Extraction

NOTE: The extraction instructions below are for use when testing individual specimens ONLY. When  pooling specimens, refer to Appendix B for modified extraction instructions.

Performance of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel is dependent upon the amount

and quality of template RNA purified from human specimens. The following commercially available  RNA extraction kits and procedures have been qualified and validated for recovery and purity of RNA

for use with the panel:

Qiagen QIAamp® DSP Viral RNA Mini Kit or QIAamp® Viral RNA Mini Kit

Recommendation(s): Utilize 100 μL of sample and elute with 100 μL of buffer or utilize 140 μL of

sample and elute with 140 μL of buffer.

14

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Qiagen EZ1 Advanced XL

Kit: Qiagen EZ1 DSP Virus Kit and Buffer AVL (supplied separately) for offboard lysis

Card: EZ1 Advanced XL DSP Virus Card

Recommendation(s): Add 120 μL of sample to 280 μL of pre-aliquoted Buffer AVL (total input sample  volume is 400 μL). Proceed with the extraction on the EZ1 Advanced XL. Elution volume is 120 μL.

Kit: Qiagen EZ1 Virus Mini Kit v2.0 and Buffer AVL (supplied separately) for offboard lysis  Card: EZ1 Advanced XL Virus Card v2.0

Recommendation(s): Add 120 μL of sample to 280 μL of pre-aliquoted Buffer AVL (total input sample  volume is 400 μL). Proceed with the extraction on the EZ1 Advanced XL. Elution volume is 120 μL.

Roche MagNA Pure 96

Kit: Roche MagNA Pure 96 DNA and Viral NA Small Volume Kit

Protocol: Viral NA Plasma Ext LysExt Lys SV 4.0 Protocol or Viral NA Plasma Ext Lys SV Protocol

Recommendation(s): Add 100 μL of sample to 350 μL of pre-aliquoted External Lysis Buffer (supplied  separately) (total input sample volume is 450 μL). Proceed with the extraction on the MagNA Pure 96.

(Internal Control = None). Elution volume is 100 μL.

Roche MagNA Pure 24

Kit: Roche MagNA Pure 24 Total NA Isolation Kit

Protocol: Pathogen 1000 2.0 Protocol

Recommendation(s): Add 100 µL of sample to 400 µL of pre-aliquoted External Lysis Buffer (supplied  separately) (total input sample volume is 500 µL). Proceed with the extraction on the MagNA Pure 24.

(Internal Control = None). Elution volume is 100 µL.

Promega Maxwell® RSC 48 and Maxwell® CSC 48

Kit: Promega Maxwell® Viral Total Nucleic Acid Purification Kit

Protocol: Viral Total Nucleic Acid

Recommendation(s): Add 120 µL of sample to 330 µL of pre-aliquoted External Lysis Buffer (300 µL  Lysis Buffer plus 30 µL Proteinase K; supplied within the kit) (total input volume is 450 µL). Proceed  with the extraction on the Maxwell® RSC 48. Elution volume is 75 µL.

Equivalence and performance of the following extraction platforms were demonstrated with the CDC

Human Influenza Virus Real-Time RT-PCR Diagnostic Panel (K190302) and based on those data are

acceptable for use with the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel.

QIAGEN QIAcube

Kit: QIAGEN QIAamp® DSP Viral RNA Mini Kit or QIAamp® Viral RNA Mini Kit

Recommendations: Utilize 140 μL of sample and elute with 100 μL of buffer.

15

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Roche MagNA Pure LC

Kit: Roche MagNA Pure Total Nucleic Acid Kit

Protocol: Total NA External_lysis

Recommendation(s): Add 100 μL of sample to 300 μL of pre-aliquoted TNA isolation kit lysis buffer  (total input sample volume is 400 μL). Elution volume is 100 μL.

Roche MagNA Pure Compact

Kit: Roche MagNA Pure Nucleic Acid Isolation Kit I

Protocol: Total_NA_Plasma100_400

Recommendation(s): Add 100 μL of sample to 300 μL of pre-aliquoted TNA isolation kit lysis buffer  (total input sample volume is 400 μL). Elution volume is 100 μL.

bioMérieux NucliSENS® easyMAG® Instrument

Protocol: General protocol (not for blood) using “Off-board Lysis” reagent settings.  Recommendation(s): Add 100 μL of sample to 1000 μL of pre-aliquoted easyMAG lysis buffer (total  input sample volume is 1100 μL). Incubate for 10 minutes at room temperature.

Elution volume is 100 μL.

bioMérieux EMAG® Instrument

Protocol: Custom protocol: CDC Flu V1 using “Off-board Lysis” reagent settings.  Recommendation(s): Add 100 μL of samples to 2000 μL of pre-aliquoted easyMAG lysis buffer (total  input sample volume is 2100 μL). Incubate for 10 minutes at room temperature. Elution volume is 100

μL. The custom protocol, CDC Flu V1, is programmed on the bioMérieux EMAG® instrument with the

assistance of a bioMérieux service representative. Installation verification is documented at the time of

installation. Laboratories are recommended to retain a record of the step-by-step verification of the

bioMérieux custom protocol installation procedure.

Manufacturer’s recommended procedures (except as noted in recommendations above) are to be

followed for sample extraction. HSC must be included in each extraction batch.

Disclaimer: Names of vendors or manufacturers are provided as examples of suitable product sources. Inclusion

does not imply endorsement by the Centers for Disease Control and Prevention.

Assay Set Up

Reaction Master Mix and Plate Set Up

Note: Plate set-up configuration can vary with the number of specimens and workday  organization. NTCs and nCoVPCs must be included in each run.

1) In the reagent set-up room clean hood, place rRT-PCR buffer, enzyme, and primer/probes on

ice or cold-block. Keep cold during preparation and use.

2) Mix buffer, enzyme, and primer/probes by inversion 5 times.

3) Centrifuge reagents and primers/probes for 5 seconds to collect contents at the bottom of

the tube, and then place the tube in a cold rack.

4) Label one 1.5 mL microcentrifuge tube for each primer/probe set.

16

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

5) Determine the number of reactions (N) to set up per assay. It is necessary to make excess  reaction mix for the NTC, nCoVPC, HSC (if included in the rRT-PCR run), and RP reactions and  for pipetting error. Use the following guide to determine N:

• If number of samples (n) including controls equals 1 through 14, then N = n + 1  • If number of samples (n) including controls is 15 or greater, then N = n + 2

6) For each primer/probe set, calculate the amount of each reagent to be added for each

reaction mixture (N = # of reactions).

17

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix

| Step # | Reagent | Vol. of Reagent Added per Reaction |
| --- | --- | --- |
| 1 | Nuclease-free Water | N x 8.5 µL |
| 2 | Combined Primer/Probe Mix | N x 1.5 µL |
| 3 | TaqPathTM 1-Step RT-qPCR Master Mix (4x) | N x 5.0 µL |
|  | Total Volume | N x 15.0 µL |

Promega GoTaq® Probe 1- Step RT-qPCR System

| Step # | Reagent | Vol. of Reagent Added per Reaction |
| --- | --- | --- |
| 1 | Nuclease-free Water | N x 3.1 µL |
| 2 | Combined Primer/Probe Mix | N x 1.5 µL |
| 3 | GoTaq Probe qPCR Master Mix with dUTP | N x 10.0 µL |
| 4 | Go Script RT Mix for 1-Step RT-qPCR | N x 0.4 µL |
|  | Total Volume | N x 15.0 µL |

Quantabio qScript XLT One-Step RT-qPCR ToughMix

| Step # | Reagent | Vol. of Reagent Added per Reaction |
| --- | --- | --- |
| 1 | Nuclease-free Water | N x 3.5 µL |
| 2 | Combined Primer/Probe Mix | N x 1.5 µL |
| 3 | qScript XLT One-Step RT-qPCR ToughMix(2X) | N x 10.0 µL |
|  | Total Volume | N x 15.0 µL |

Quantabio UltraPlex 1-Step ToughMix (4X)

| Step # | Reagent | Vol. of Reagent Added per Reaction |
| --- | --- | --- |
| 1 | Nuclease-free Water | N x 8.5 µL |
| 2 | Combined Primer/Probe Mix | N x 1.5 µL |
| 3 | UltraPlex 1-Step ToughMix (4X) | N x 5.0 µL |
|  | Total Volume | N x 15.0 µL |

18

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

7) Dispense reagents into each respective labeled 1.5 mL microcentrifuge tube. After addition of  the reagents, mix reaction mixtures by pipetting up and down. Do not vortex.

8) Centrifuge for 5 seconds to collect contents at the bottom of the tube, and then place the  tube in a cold rack.

9) Set up reaction strip tubes or plates in a 96-well cooler rack.

10) Dispense 15 µL of each master mix into the appropriate wells going across the row as shown  below (Figure 1):

Figure 1: Example of Reaction Master Mix Plate Set-Up

| A | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | N1 | N1 | N1 | N1 | N1 | N1 | N1 | N1 | N1 | N1 | N1 | N1 |
| B | N2 | N2 | N2 | N2 | N2 | N2 | N2 | N2 | N2 | N2 | N2 | N2 |
| C | RP | RP | RP | RP | RP | RP | RP | RP | RP | RP | RP | RP |
| D |  |  |  |  |  |  |  |  |  |  |  |  |
| E |  |  |  |  |  |  |  |  |  |  |  |  |
| F |  |  |  |  |  |  |  |  |  |  |  |  |
| G |  |  |  |  |  |  |  |  |  |  |  |  |
| H |  |  |  |  |  |  |  |  |  |  |  |  |

11) Prior to moving to the nucleic acid handling area, prepare the No Template Control (NTC)

reactions for column #1 in the assay preparation area.

12) Pipette 5 µL of nuclease-free water into the NTC sample wells (Figure 2, column 1). Securely

cap NTC wells before proceeding.

13) Cover the entire reaction plate and move the reaction plate to the specimen nucleic acid  handling area.

Nucleic Acid Template Addition

1) Gently vortex nucleic acid sample tubes for approximately 5 seconds.

2) Centrifuge for 5 seconds to collect contents at the bottom of the tube.

3) After centrifugation, place extracted nucleic acid sample tubes in the cold rack.

19

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

4) Samples should be added to columns 2-11 (column 1 and 12 are for controls) to the specific  assay that is being tested as illustrated in Figure 2. Carefully pipette 5.0 µL of the first sample  into all the wells labeled for that sample (i.e. Sample “S1” down column #2). Keep other  sample wells covered during addition. Change tips after each addition.

5) Securely cap the column to which the sample has been added to prevent cross contamination

and to ensure sample tracking.

6) Change gloves often and when necessary to avoid contamination.

7) Repeat steps #4 and #5 for the remaining samples.

8) If necessary, add 5 µL of Human Specimen Control (HSC) extracted sample to the HSC wells

(Figure 2, column 11). Securely cap wells after addition. NOTE: Per CLIA regulations, HSC

must be tested at least once per day.

9) Cover the entire reaction plate and move the reaction plate to the positive template control

handling area.

Assay Control Addition

1) Pipette 5 µL of nCoVPC RNA to the sample wells of column 12 (Figure 2). Securely cap wells  after addition of the control RNA.

NOTE: If using 8-tube strips, label the TAB of each strip to indicate sample position. DO NOT

LABEL THE TOPS OF THE REACTION TUBES!

2) Briefly centrifuge reaction tube strips for 10-15 seconds. After centrifugation return to cold  rack.

NOTE: If using 96-well plates, centrifuge plates for 30 seconds at 500 x g, 4°C.

Figure 2. 2019-nCoV rRT-PCR Diagnostic Panel: Example of Sample and Control Set-up

|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11a | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | NTC | S1 | S2 | S3 | S4 | S5 | S6 | S7 | S8 | S9 | S10 | nCoV PC |
| B | NTC | S1 | S2 | S3 | S4 | S5 | S6 | S7 | S8 | S9 | S10 | nCoV PC |
| C | NTC | S1 | S2 | S3 | S4 | S5 | S6 | S7 | S8 | S9 | S10 | nCoV PC |
| D |  |  |  |  |  |  |  |  |  |  |  |  |
| E |  |  |  |  |  |  |  |  |  |  |  |  |
| F |  |  |  |  |  |  |  |  |  |  |  |  |
| G |  |  |  |  |  |  |  |  |  |  |  |  |
| H |  |  |  |  |  |  |  |  |  |  |  |  |

aReplace the sample in this column with extracted HSC if necessary

20

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Create a Run Template on the Applied Biosystems 7500 Fast Dx Real-time PCR  Instrument (Required if no template exists)

If the template already exists on your instrument, please proceed to the RUNNING A TEST section.

1) Launch the Applied Biosystems 7500 Fast Dx Real-time PCR Instrument by double clicking on the

Applied Biosystems 7500 Fast Dx System icon on the desktop.

2) A new window should appear, select Create New Document from the menu.

Figure 3. New Document Wizard Window

Make sure to change

Run Mode to

STANDARD 7500

3) The New Document Wizard screen in Figure 3 will appear. Select:

a. Assay: Standard Curve (Absolute Quantitation)

b. Container: 96-Well Clear

c. Template: Blank Document

d. Run Mode: Standard 7500

e. Operator: Your Name

f. Comments: SDS v1.4

g. Plate Name: Your Choice

4) After making selections click Next at the bottom of the window.

21

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 4. Creating New Detectors

NOTE: ROX is the default passive reference. This will be changed to “none” in step 12.

5) After selecting next, the Select Detectors screen (Figure 4) will appear.

6) Click the New Detector button (see Figure 4).

7) The New Detector window will appear (Figure 5). A new detector will need to be defined for  each primer and probe set. Creating these detectors will enable you to analyze each primer and

probe set individually at the end of the reaction.

22

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 5. New Detector Window

8) Start by creating the N1 Detector. Include the following:

a. Name: N1

b. Description: leave blank

c. Reporter Dye: FAM

d. Quencher Dye: (none)

e. Color: to change the color of the detector indicator do the following:

⇒ Click on the color square to reveal the color chart

⇒ Select a color by clicking on one of the squares

⇒ After selecting a color click OK to return to the New Detector screen

f. Click the OK button of the New Detector screen to return to the screen shown in Figure 4.  9) Repeat step 6-8 for each target in the panel.

| Name | Reporter Dye | Quencher Dye |
| --- | --- | --- |
| N1 | FAM | (none) |
| N2 | FAM | (none) |
| RP | FAM | (none) |

23

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

10) After each Detector is added, the Detector Name, Description, Reporter and Quencher fields  will become populated in the Select Detectors screen (Figure 6).

11) Before proceeding, the newly created detectors must be added to the document. To add the

new detectors to the document, click ADD (see Figure 6). Detector names will appear on the  right-hand side of the Select Detectors window (Figure 6).

Figure 6. Adding New Detectors to Document

12) Once all detectors have been added, select (none) for Passive Reference at the top right-hand

drop-down menu (Figure 7).

Figure 7. Select Passive Reference

Passive reference should be set to “(none)” as described above.

24

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

13)Click Next at the bottom of the Select Detectors window to proceed to the Set Up Sample Plate  window (Figure 8).

14)In the Set Up Sample Plate window (Figure 8), use your mouse to select row A from the lower

portion of the window, in the spreadsheet (see Figure 8).

15)In the top portion of the window, select detector N1. A check will appear next to the detector

you have selected (Figure 8). You will also notice the row in the spreadsheet will be populated  with a colored “U” icon to indicate which detector you’ve selected.

16)Repeat step 14-15 for each detector that will be used in the assay.

Figure 8. Sample Plate Set-up

17) Select Finish after detectors have been assigned to their respective rows. (Figure 9).

Figure 9. Finished Plate Set-up

25

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

18) After clicking “Finish”, there will be a brief pause allowing the Applied Biosystems 7500 Fast Dx

to initialize. This initialization is followed by a clicking noise. Note: The machine must be turned  on for initialization.

19) After initialization, the Plate tab of the Setup (Figure 10) will appear.

20)Each well of the plate should contain colored U icons that correspond with the detector labels

that were previously chosen. To confirm detector assignments, select Tools from the file menu,  then select Detector Manager.

Figure 10. Plate Set-up Window

21) The Detector Manager window will appear (Figure 11).

Figure 11. Detector Manager Window

26

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

22)Confirm all detectors are included and that each target has a Reporter set to FAM and the

Quencher is set to (none).

23)If all detectors are present, select Done. The detector information has been created and  assigned to wells on the plate.

Defining the Instrument Settings

1) After detectors have been created and assigned, proceed to instrument set up.

2) Select the Instrument tab to define thermal cycling conditions.

3) Modify the thermal cycling conditions as follows (Figure 12):

Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix, CG

a. In Stage 1, Set to 2 min at 25°C; 1 Rep.

b. In Stage 2, Set to 15 min at 50°C; 1 Rep.

c. In Stage 3, Set to 2 min at 95°C, 1 Rep.

d. In Stage 4, Step 1 set to 3 sec at 95°C.

e. In Stage 4, Step 2 set to 30 sec at 55.0°C.

f. In Stage 4, Reps should be set to 45.

g. Under Settings (Figure 12), bottom left-hand box, change volume to 20 µL.

h. Under Settings, Run Mode selection should be Standard 7500.

i. Step 2 of Stage 4 should be highlighted in yellow to indicate data collection (see Figure  12).

OR

Quantabio qScriptTM XLT One-Step RT-qPCR ToughMix or UltraPlex 1-Step ToughMix (4X)

a. In Stage 1, Set to 10 min at 50°C; 1 Rep.

b. In Stage 2, Set to 3 min at 95°C, 1 Rep.

c. In Stage 3, Step 1 set to 3 sec at 95°C.

d. In Stage 3, Step 2 set to 30 sec at 55.0°C.

e. In Stage 3, Reps should be set to 45.

f. Under Settings (Figure 12), bottom left-hand box, change volume to 20 µL.  g. Under Settings, Run Mode selection should be Standard 7500.

h. Step 2 of Stage 3 should be highlighted in yellow to indicate data collection (see Figure  12).

OR

Promega GoTaq® Probe 1-Step RT-qPCR System

a. In Stage 1, Set to 15 min at 45°C; 1 Rep.

b. In Stage 2, Set to 2 min at 95°C, 1 Rep.

27

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

c. In Stage 3, Step 1 set to 3 sec at 95°C.  d. In Stage 3, Step 2 set to 30 sec at 55.0°C.  e. In Stage 3, Reps should be set to 45.

f. Under Settings (Figure 12), bottom left-hand box, change volume to 20 µL.  g. Under Settings, Run Mode selection should be Standard 7500.

h. Step 2 of Stage 3 should be highlighted in yellow to indicate data collection (see Figure  12).

Figure 12. Instrument Window

4) After making changes to the Instrument tab, the template file is ready to be saved. To save the

template, select File from the top menu, then select Save As. Since the enzyme options have

different instrument settings, it is recommended that the template be saved with a name

indicating the enzyme option.

5) Save the template as 2019-nCoV Dx Panel TaqPath or 2019-nCoV Dx Panel Quanta or 2019-

nCoV Dx Panel Promega as appropriate in the desktop folder labeled “ABI Run Templates” (you  must create this folder). Save as type should be SDS Templates (*.sdt) (Figure 13).

28

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 13. Saving Template

Running a Test

1) Turn on the ABI 7500 Fast Dx Real-Time PCR Instrument.

2) Launch the Applied Biosystems 7500 Fast Dx Real-time PCR System by double clicking on the  7500 Fast Dx System icon on the desktop.

3) A new window should appear, select Open Existing Document from the menu.  4) Navigate to select your ABI Run Template folder from the desktop.

5) Double click on the appropriate template file (2019-nCoV Dx Panel TaqPath or 2019-nCoV Dx

Panel Quanta or 2019-nCoV Dx Panel Promega)

6) There will be a brief pause allowing the Applied Biosystems 7500 Fast Dx Real-Time PCR  Instrument to initialize. This initialization is followed by a clicking noise. Note: The machine must  be turned on for initialization.

29

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 14. Plate Set-up Window

7) After the instrument initializes, a plate map will appear (Figure 14). The detectors and controls

should already be labeled as they were assigned in the original template.

8) Click the Well Inspector icon from the top menu.

9) Highlight specimen wells of interest on the plate map.

10) Type sample identifiers to Sample Name box in the Well Inspector window (Figure 15).

30

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 15. Labeling Wells

11) Repeat steps 9-10 until all sample identifiers are added to the plate setup.

12) Once all specimen and control identifiers are added click the Close button on the Well Inspector  window to return to the Plate set up tab.

13) Click the Instrument tab at the upper left corner.

14) The reaction conditions, volumes, and type of 7500 reaction should already be loaded (Figure  16).

31

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 16. Instrument Settings

15) Ensure settings are correct (refer to the Defining Instrument Settings).

16) Before proceeding, the run file must be saved; from the main menu, select File, then Save As.  Save in appropriate run folder designation.

17) Load the plate into the plate holder in the instrument. Ensure that the plate is properly aligned

in the holder.

18) Once the run file is saved, click the Start button. Note: The run should take approximately 1 hour  and 20 minutes to complete.

32

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Data Analysis

1) After the run has completed, select the Results tab at the upper left corner of the software.  2) Select the Amplification Plot tab to view the raw data (Figure 17).

Figure 17. Amplification Plot Window

b cd

e

a

c

c

3) Start by highlighting all the samples from the run; to do this, click on the upper left-hand box (a)

of the sample wells (Figure 17). All the growth curves should appear on the graph.

4) On the right-hand side of the window (b), the Data drop down selection should be set to Delta  Rn vs. Cycle.

5) Select N1 from (c), the Detector drop down menu, using the downward arrow.

a. Please note that each detector is analyzed individually to reflect different

performance profiles of each primer and probe set.

6) In the Line Color drop down (d), Detector Color should be selected.

7) Under Analysis Settings select Manual Ct (e).

b. Do not change the Manual Baseline default numbers.

8) Using the mouse, click and drag the red threshold line until it lies within the exponential phase  of the fluorescence curves and above any background signal (Figure 18).

33

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Figure 18. Amplification Plot

Exponential

PCR Phase

Background noise

Threshold adjusted

to fall within the

PCR exponential  phase.

9) Click the Analyze button in the lower right corner of the window. The red threshold line will turn

to green, indicating the data has been analyzed.

10) Repeat steps 5-9 to analyze results generated for each set of markers (N1, N2, RP).  11) Save analysis file by selecting File then Save As from the main menu.

12) After completing analysis for each of the markers, select the Report tab above the graph to

display the Ct values (Figure 19). To filter report by sample name in ascending or descending

order, simply click on Sample Name in the table.

Figure 19. Report

34

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Interpretation of Results and Reporting

Extraction and Positive Control Results and Interpretation

No Template Control (NTC)

The NTC consists of using nuclease-free water in the rRT-PCR reactions instead of RNA. The NTC

reactions for all primer and probe sets should not exhibit fluorescence growth curves that cross the  threshold line. If any of the NTC reactions exhibit a growth curve that crosses the cycle threshold, sample  contamination may have occurred. Invalidate the run and repeat the assay with strict adherence to the

guidelines.

2019-nCoV Positive Control (nCoVPC)

The nCoVPC consists of in vitro transcribed RNA. The nCoVPC will yield a positive result with the

following primer and probe sets: N1, N2, and RP.

Human Specimen Control (HSC) (Extraction Control)

When HSC is run with the CDC 2019-nCoV rRT-PCR Diagnostic Panel (see previous section on Assay Set

Up), the HSC is used as a nucleic acid extraction procedural control to demonstrate successful recovery

of nucleic acid as well as extraction reagent integrity. The HSC control consists of noninfectious cultured

human cell (A549) material. Purified nucleic acid from the HSC should yield a positive result with the RP  primer and probe set and negative results with all 2019-nCoV markers.

35

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Expected Performance of Controls Included in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

| Control Type | External Control Name | Used to Monitor | 2019 nCoV_N1 | 2019 nCoV_N2 | RP | Expected Ct Values |
| --- | --- | --- | --- | --- | --- | --- |
| Positive | nCoVPC | Substantial reagent failure including primer and probe integrity | + | + | + | < 40.00 Ct |
| Negative | NTC | Reagent and/or environmental contamination | - | - | - | None detected |
| Extraction | HSC | Failure in lysis and extraction procedure, potential contamination during extraction | - | - | + | < 40.00 Ct |

If any of the above controls do not exhibit the expected performance as described, the assay may have  been set up and/or executed improperly, or reagent or equipment malfunction could have occurred.  Invalidate the run and re-test.

RNase P (Extraction Control)

⮚ All clinical samples should exhibit fluorescence growth curves in the RNase P reaction that cross  the threshold line within 40.00 cycles (< 40.00 Ct), thus indicating the presence of the human

RNase P gene. Failure to detect RNase P in any clinical specimens may indicate:

− Improper extraction of nucleic acid from clinical materials resulting in loss of RNA and/or

RNA degradation.

− Absence of sufficient human cellular material due to poor collection or loss of specimen  integrity.

− Improper assay set up and execution.

− Reagent or equipment malfunction.

⮚ If the RP assay does not produce a positive result for human clinical specimens, interpret as

follows:

− If the 2019-nCoV N1 and N2 are positive even in the absence of a positive RP, the result

should be considered valid. It is possible, that some samples may fail to exhibit RNase P

growth curves due to low cell numbers in the original clinical sample. A negative RP signal  does not preclude the presence of 2019-nCoV virus RNA in a clinical specimen.

− If all 2019-nCoV markers AND RNase P are negative for the specimen, the result should be  considered invalid for the specimen. If residual specimen is available, repeat the extraction

36

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

procedure and repeat the test. If all markers remain negative after re-test, report the  results as invalid and a new specimen should be collected if possible.

2019-nCoV Markers (N1 and N2)

• When all controls exhibit the expected performance, a specimen is considered negative if all  2019-nCoV marker (N1, N2) cycle threshold growth curves DO NOT cross the threshold line  within 40.00 cycles (< 40.00 Ct) AND the RNase P growth curve DOES cross the threshold line  within 40.00 cycles (< 40.00 Ct).

• When all controls exhibit the expected performance, a specimen is considered positive for 2019- nCoV if all 2019-nCoV marker (N1, N2) cycle threshold growth curves cross the threshold line

within 40.00 cycles (< 40.00 Ct). The RNase P may or may not be positive as described above, but  the 2019-nCoV result is still valid.

• When all controls exhibit the expected performance and the growth curves for the 2019-nCoV

markers (N1, N2) AND the RNase P marker DO NOT cross the cycle threshold growth curve within

40.00 cycles (< 40.00 Ct), the result is invalid. The extracted RNA from the specimen should be re tested. If residual RNA is not available, re-extract RNA from residual specimen and re-test. If the

re-tested sample is negative for all markers and RNase P, the result is invalid and collection of a  new specimen from the patient should be considered.

• When all controls exhibit the expected performance and the cycle threshold growth curve for any  one marker (N1 or N2, but not both markers) crosses the threshold line within 40.00 cycles (<

40.00 Ct) the result is inconclusive. The extracted RNA should be retested. If residual RNA is not

available, re-extract RNA from residual specimen and re-test. If the same result is obtained,  report the inconclusive result. Consult with your state public health laboratory or CDC, as

appropriate, to request guidance and/or to coordinate transfer of the specimen for additional  analysis.

• If HSC is positive for N1 or N2, then contamination may have occurred during extraction or

sample processing. Invalidate all results for specimens extracted alongside the HSC. Re-extract  specimens and HSC and re-test.

37

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

2019-nCoV rRT-PCR Diagnostic Panel Results Interpretation Guide

The table below lists the expected results for the 2019-nCoV rRT-PCR Diagnostic Panel. If a laboratory  obtains unexpected results for assay controls or if inconclusive or invalid results are obtained and cannot

be resolved through the recommended re-testing, please contact CDC for consultation and possible  specimen referral. See pages 13 and 53 for referral and contact information.

| 2019 nCoV_N1 | 2019 nCoV_N2 | RP | Result Interpretationa | Report | Actions |
| --- | --- | --- | --- | --- | --- |
| + | + | ± | 2019-nCoV detected | Positive 2019-nCoV | Report results to CDC and sender. |
| If only one of the two targets is positive |  | ± | Inconclusive Result | Inconclusive | Repeat testing of nucleic acid and/or re-extract and repeat rRT-PCR. If the repeated result remains inconclusive, contact your State Public Health Laboratory or CDC for instructions for transfer of the specimen or further guidance. |
| - | - | + | 2019-nCoV not detected | Not Detected | Report results to sender. Consider testing for other respiratory viruses.b |
| - | - | - | Invalid Result | Invalid | Repeat extraction and rRT-PCR. If the repeated result remains invalid, consider collecting a new specimen from the patient. |

aLaboratories should report their diagnostic result as appropriate and in compliance with their specific reporting  system.

bOptimum specimen types and timing for peak viral levels during infections caused by 2019-nCoV have not been determined. Collection of multiple specimens from the same patient may be necessary to detect the virus. The

possibility of a false negative result should especially be considered if the patient’s recent exposures or clinical

presentation suggest that 2019-nCoV infection is possible, and diagnostic tests for other causes of illness (e.g., other respiratory illness) are negative. If 2019-nCoV infection is still suspected, re-testing should be considered

in consultation with public health authorities.

38

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Quality Control

• Quality control requirements must be performed in conformance with local, state, and federal  regulations or accreditation requirements and the user’s laboratory’s standard quality control  procedures. For further guidance on appropriate quality control practices, refer to 42 CFR  493.1256.

• Quality control procedures are intended to monitor reagent and assay performance.  • Test all positive controls prior to running diagnostic samples with each new kit lot to ensure all

reagents and kit components are working properly.

• Good laboratory practice (cGLP) recommends including a positive extraction control in each  nucleic acid isolation batch.

• Although HSC is not included with the 2019-nCov rRT-PCR Diagnostic Panel, the HSC extraction  control must proceed through nucleic acid isolation per batch of specimens to be tested.  • Always include a negative template control (NTC) and the appropriate positive control (nCoVPC)  in each amplification and detection run. All clinical samples should be tested for human RNase P

gene to control for specimen quality and extraction.

Limitations

• All users, analysts, and any person reporting diagnostic results should be trained to perform this  procedure by a competent instructor. They should demonstrate their ability to perform the test  and interpret the results prior to performing the assay independently.

• Performance of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel has only been established  in upper and lower respiratory specimens (such as nasopharyngeal or oropharyngeal swabs,  sputum, lower respiratory tract aspirates, bronchoalveolar lavage, and nasopharyngeal

wash/aspirate or nasal aspirate).

• Negative results do not preclude 2019-nCoV infection and should not be used as the sole basis

for treatment or other patient management decisions. Optimum specimen types and timing for  peak viral levels during infections caused by 2019-nCoV have not been determined. Collection of  multiple specimens (types and time points) from the same patient may be necessary to detect  the virus.

• A false-negative result may occur if a specimen is improperly collected, transported or handled.

False-negative results may also occur if amplification inhibitors are present in the specimen or if  inadequate numbers of organisms are present in the specimen.

• Positive and negative predictive values are highly dependent on prevalence. False-negative test

results are more likely when prevalence of disease is high. False-positive test results are more

likely when prevalence is moderate to low.

• Do not use any reagent past the expiration date.

• If the virus mutates in the rRT-PCR target region, 2019-nCoV may not be detected or may be

detected less predictably. Inhibitors or other types of interference may produce a false-negative

result. An interference study evaluating the effect of common cold medications was not  performed.

• Test performance can be affected because the epidemiology and clinical spectrum of infection  caused by 2019-nCoV is not fully known. For example, clinicians and laboratories may not know

39

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

the optimum types of specimens to collect, and, during the course of infection, when these  specimens are most likely to contain levels of viral RNA that can be readily detected.  • Detection of viral RNA may not indicate the presence of infectious virus or that 2019-nCoV is the  causative agent for clinical symptoms.

• The performance of this test has not been established for monitoring treatment of 2019-nCoV

infection.

• The performance of this test has not been established for screening of blood or blood products

for the presence of 2019-nCoV.

• This test cannot rule out diseases caused by other bacterial or viral pathogens.

Conditions of Authorization for the Laboratory

The CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel Letter of Authorization, along with the  authorized Fact Sheet for Healthcare Providers, the authorized Fact Sheet for Patients, and  authorized labeling are available on the FDA website: https://www.fda.gov/medical

devices/coronavirus-disease-2019-covid-19-emergency-use-authorizations-medical-devices/vitro diagnostics-euas.

However, to assist clinical laboratories using the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

(“your product” in the conditions below), the relevant Conditions of Authorization are listed below:

• Authorized laboratories using your product will include with test result reports, all authorized

Fact Sheets available on the CDC website. Under exigent circumstances, other appropriate  methods for disseminating these Fact Sheets may be used, which may include mass media.

• Authorized laboratories using your product will use your product as outlined in the authorized

labeling available on the CDC website. Deviations from the authorized procedures, including the  authorized RT-PCR instruments, authorized extraction methods, authorized clinical specimen  types, authorized control materials, authorized other ancillary reagents and authorized materials  required to use your product are not permitted under this authorization.

• Authorized laboratories that receive the commercially manufactured and distributed primer and

probe sets identified as acceptable on the CDC website for use with your product, and are not

able to obtain the authorized Human Specimen Control and authorized Positive Control for 2019-

nCoV (NCoVPC) materials described in your product’s authorized labeling, may use appropriate  materials identified as acceptable materials on the CDC website for use with your product.

• Authorized laboratories that receive your product will notify the relevant public health

authorities of their intent to run your product prior to initiating testing.

• Authorized laboratories using your product will have a process in place for reporting test results

to healthcare providers and relevant public health authorities, as appropriate.

• Authorized laboratories will collect information on the performance of your product and report  to DMD/OHT7-OIR/OPEQ/CDRH (via email: CDRH-EUA-Reporting@fda.hhs.gov) and CDC  (respvirus@cdc.gov) any suspected occurrence of false positive or false negative results and

40

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

significant deviations from the established performance characteristics of the test of which they  become aware.

• Authorized laboratories using specimen pooling strategies when testing patient specimens with  your product will include with negative test result reports for specific patients whose specimen(s)  were the subject of pooling, a notice that pooling was used during testing and that “Patient

specimens with low viral loads may not be detected in sample pools due to the decreased

sensitivity of pooled testing.”

• Authorized laboratories implementing pooling strategies for testing patient specimens must use  the “Implementation and Monitoring of Pooled Specimen Testing” available in the authorized  labeling to evaluate the appropriateness of continuing to use such strategies based on the  recommendations in the protocol.

• Authorized laboratories will keep records of specimen pooling strategies implemented including  type of strategy, date implemented, and quantities tested, and test result data generated as part  of the Protocol for Monitoring of Specimen Pooling Testing Strategies. For the first 12 months  from the date of their creation, such records will be made available to FDA within 48 business  hours (2 business days) for inspection upon request, and will be made available within a

reasonable time after 12 months from the date of their creation.

• Authorized laboratories will report adverse events, including problems with your products

performance or results, to MedWatch by submitting the online FDA Form 3500

(https://www.accessdata.fda.gov/scripts/medwatch/index.cfm?action=reporting.home) or by  calling 1-800-FDA-1088.

• All laboratory personnel using the test must be appropriately trained in RT-PCR techniques and

use appropriate laboratory and personal protective equipment when handling this kit and use

the test in accordance with the authorized labeling.

• CDC, IRR, manufacturers and distributors of commercial materials identified as acceptable on the

CDC website, and authorized laboratories using your product will ensure that any records

associated with this EUA are maintained until otherwise notified by FDA. Such records will be

made available to FDA for inspection upon request.

41

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Performance Characteristics

Analytical Performance:

Limit of Detection (LoD):

LoD studies determine the lowest detectable concentration of 2019-nCoV at which approximately 95%

of all (true positive) replicates test positive. The LoD was determined by limiting dilution studies using

characterized samples.

The analytical sensitivity of the rRT-PCR assays contained in the CDC 2019 Novel Coronavirus (2019- nCoV) Real-Time RT-PCR Diagnostic Panel were determined in Limit of Detection studies. Since no  quantified virus isolates of the 2019-nCoV were available for CDC use at the time the test was developed

and this study conducted, assays designed for detection of the 2019-nCoV RNA were tested with  characterized stocks of in vitro transcribed full length RNA (N gene; GenBank accession: MN908947.2) of  known titer (RNA copies/µL) spiked into a diluent consisting of a suspension of human A549 cells and

viral transport medium (VTM) to mimic clinical specimen. Samples were extracted using the QIAGEN EZ1  Advanced XL instrument and EZ1 DSP Virus Kit (Cat# 62724) and manually with the QIAGEN DSP Viral

RNA Mini Kit (Cat# 61904). Real-Time RT-PCR assays were performed using the Thermo Fisher Scientific  TaqPath™ 1-Step RT-qPCR Master Mix, CG (Cat# A15299) on the Applied Biosystems™ 7500 Fast Dx Real

Time PCR Instrument according to the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel instructions  for use.

A preliminary LoD for each assay was determined testing triplicate samples of RNA purified using each  extraction method. The approximate LoD was identified by extracting and testing 10-fold serial dilutions  of characterized stocks of in vitro transcribed full-length RNA. A confirmation of the LoD was determined  using 3-fold serial dilution RNA samples with 20 extracted replicates. The LoD was determined as the  lowest concentration where ≥ 95% (19/20) of the replicates were positive.

Table 4. Limit of Detection Confirmation of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel  with QIAGEN EZ1 DSP

| Targets | 2019-nCoV_N1 |  |  | 2019-nCoV_N2 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| RNA Concentration1 | 10 0.5 | 10 0.0 | 10 -0.5 | 10 0.5 | 10 0.0 | 10 -0.5 |
| Positives/Total | 20/20 | 19/20 | 13/20 | 20/20 | 17/20 | 9/20 |
| Mean Ct2 | 32.5 | 35.4 | NA | 35.8 | NA | NA |
| Standard Deviation (Ct) | 0.5 | 0.8 | NA | 1.3 | NA | NA |

1 Concentration is presented in RNA copies/µL

2 Mean Ct reported for dilutions that are ≥ 95% positive. Calculations only include positive results.  NA not applicable

42

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table 5. Limit of Detection Confirmation CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel with  QIAGEN QIAmp DSP Viral RNA Mini Kit

| Targets | 2019-nCoV_N1 |  |  | 2019-nCoV_N2 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| RNA Concentration1 | 10 0.5 | 10 0.0 | 10 -0.5 | 10 0.5 | 10 0.0 | 10 -0.5 | 10 -1.0 |
| Positives/Total | 20/20 | 20/20 | 6/20 | 20/20 | 20/20 | 20/20 | 8/20 |
| Mean Ct2 | 32.0 | 32.8 | NA | 33.0 | 35.4 | 36.2 | NA |
| Standard Deviation (Ct) | 0.7 | 0.8 | NA | 1.4 | 0.9 | 1.9 | NA |

1 Concentration is presented in RNA copies/µL

2 Mean Ct reported for dilutions that are ≥ 95% positive. Calculations only include positive results.  NA not applicable

Table 6. Limit of Detection of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

| Virus | Material | Limit of Detection (RNA copies/∝L) |  |
| --- | --- | --- | --- |
|  |  | QIAGEN EZ1 Advanced XL | QIAGEN DSP Viral RNA Mini Kit |
| 2019 Novel Coronavirus | N Gene RNA Transcript | 100.5 | 100 |

FDA Sensitivity Evaluation: The analytical sensitivity of the test will be further assessed by evaluating an  FDA-recommended reference material using an FDA developed protocol if applicable and/or when  available.

FDA SARS-CoV-2 Reference Panel Testing

The evaluation of sensitivity and MERS-CoV cross-reactivity was performed using reference material (T1),  blinded samples and a standard protocol provided by the FDA. The study included a range finding study

and a confirmatory study for LoD. Blinded sample testing was used to establish specificity and to confirm  the LoD. Samples were extracted using the QIAGEN EZ1 Advanced XL with the QIAGEN EZ1 DSP Virus Kit.

Extracted samples were then tested using the 2019-nCoV Real-Time RT-PCR Diagnostic Panel on the

Applied BioSystems 7500 Fast Dx Real-Time PCR Instrument using the ThermoFisher TaqPath™ 1-Step

RT-qPCR Master Mix. The results are summarized in Table 7.

Table 7: Summary of LoD Confirmation Result using the FDA SARS-CoV-2 Reference Panel

| Reference Materials Provided by FDA | Specimen Type | Product LoD | Cross Reactivity |
| --- | --- | --- | --- |
| SARS-CoV-2 | NP swab | 1.8x104 NDU/mL | N/A |
| MERS-CoV |  | N/A | ND |

NDU/mL = RNA NAAT detectable units/mL

N/A: Not applicable

ND: Not detected

43

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

In Silico Analysis of Primer and Probe Sequences:

The oligonucleotide primer and probe sequences of the CDC 2019 nCoV Real-Time RT-PCR Diagnostic  Panel were evaluated against 31,623 sequences available in the Global Initiative on Sharing All Influenza

Data (GISAID, https://www.gisaid.org) database as of June 20, 2020, to demonstrate the predicted  inclusivity of the 2019-nCoV Real-Time RT-PCR Diagnostic Panel. Nucleotide mismatches in the

primer/probe regions with frequencies > 0.1% are shown below. With the exception of one nucleotide

mismatch with frequency > 1% (2.00%) at the third position of the N1 probe, the frequency of all

mismatches was < 1%, indicating that prevalence of the mismatches were sporadic. Only one sequence

(0.0032%) had two nucleotide mismatches in the N1 probe, and one other sequence from a different

isolate (0.0032%) had two nucleotide mismatches in the N1 reverse primer. No sequences were found to

have more than one mismatch in any N2 primer/probe region. The risk of these mismatches resulting in

a significant loss in reactivity causing a false negative result is extremely low due to the design of the  primers and probes, with melting temperatures > 60°C and with annealing temperature at 55°C that can

tolerate up to two mismatches.

Table 8. In Silico Inclusivity Analysis of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel Among

31,623 Genome Sequences Available from GISAID as of June 20, 2020

| Primer/probe | N1 probe | N1 reverse |  | N2 probe |
| --- | --- | --- | --- | --- |
| Location (5'>3') | 3 | 15 | 21 | 13 |
| Mismatch Nucleotide | C>T | G>T | T>C | C>T |
| Mismatch No. | 632 | 34 | 71 | 46 |
| Mismatch Frequency (%) | 2.00 | 0.11 | 0.22 | 0.15 |

Specificity/Exclusivity Testing: In Silico Analysis

BLASTn analysis queries of the 2019-nCoV rRT-PCR assays primers and probes were performed against

public domain nucleotide sequences. The database search parameters were as follows: 1) The nucleotide

collection consists of GenBank+EMBL+DDBJ+PDB+RefSeq sequences, but excludes EST, STS, GSS, WGS,

TSA, patent sequences as well as phase 0, 1, and 2 HTGS sequences and sequences longer than 100Mb;  2) The database is non-redundant. Identical sequences have been merged into one entry, while

preserving the accession, GI, title and taxonomy information for each entry; 3) Database was updated on  10/03/2019; 4) The search parameters automatically adjust for short input sequences and the expect  threshold is 1000; 5) The match and mismatch scores are 1 and -3, respectively; 6) The penalty to create

and extend a gap in an alignment is 5 and 2 respectively.

2019-nCoV_N1 Assay:

Probe sequence of 2019-nCoV rRT-PCR assay N1 showed high sequence homology with SARS  coronavirus and Bat SARS-like coronavirus genome. However, forward and reverse primers showed no  sequence homology with SARS coronavirus and Bat SARS-like coronavirus genome. Combining primers  and probe, there is no significant homologies with human genome, other coronaviruses or human

microflora that would predict potential false positive rRT-PCR results.

44

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

2019-nCoV_N2 Assay:

The forward primer sequence of 2019-nCoV rRT-PCR assay N2 showed high sequence homology to Bat

SARS-like coronaviruses. The reverse primer and probe sequences showed no significant homology with

human genome, other coronaviruses or human microflora. Combining primers and probe, there is no

prediction of potential false positive rRT-PCR results.

In summary, the 2019-nCoV rRT-PCR assay N1 and N2, designed for the specific detection of 2019-nCoV,  showed no significant combined homologies with human genome, other coronaviruses, or human  microflora that would predict potential false positive rRT-PCR results.

In addition to the in silico analysis, several organisms were extracted and tested with the CDC 2019-nCoV

Real-Time RT-PCR Diagnostic Panel to demonstrate analytical specificity and exclusivity. Studies were

performed with nucleic acids extracted using the QIAGEN EZ1 Advanced XL instrument and EZ1 DSP  Virus Kit. Nucleic acids were extracted from high titer preparations (typically ≥ 105 PFU/mL or ≥ 106

CFU/mL). Testing was performed using the Thermo Fisher Scientific TaqPath™ 1-Step RT-qPCR Master

Mix, CG on the Applied Biosystems™ 7500 Fast Dx Real-Time PCR instrument. The data demonstrate that  the expected results are obtained for each organism when tested with the CDC 2019-nCoV Real-Time RT PCR Diagnostic Panel.

Table 9. Specificity/Exclusivity of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

| Virus | Strain | Source | 2019- nCoV_ N1 | 2019- nCoV_ N2 | Final Result |
| --- | --- | --- | --- | --- | --- |
| Human coronavirus | 229E | Isolate | 0/3 | 0/3 | Neg. |
| Human coronavirus | OC43 | Isolate | 0/3 | 0/3 | Neg. |
| Human coronavirus | NL63 | clinical specimen | 0/3 | 0/3 | Neg. |
| Human coronavirus | HKU1 | clinical specimen | 0/3 | 0/3 | Neg. |
| MERS-coronavirus |  | Isolate | 0/3 | 0/3 | Neg. |
| SARS-coronavirus |  | Isolate | 0/3 | 0/3 | Neg. |
| bocavirus | - | clinical specimen | 0/3 | 0/3 | Neg. |
| Mycoplasma pneumoniae |  | Isolate | 0/3 | 0/3 | Neg. |
| Streptococcus |  | Isolate | 0/3 | 0/3 | Neg. |
| Influenza A(H1N1) |  | Isolate | 0/3 | 0/3 | Neg. |
| Influenza A(H3N2) |  | Isolate | 0/3 | 0/3 | Neg. |
| Influenza B |  | Isolate | 0/3 | 0/3 | Neg. |
| Human adenovirus, type 1 | Ad71 | Isolate | 0/3 | 0/3 | Neg. |
| Human metapneumovirus | - | Isolate | 0/3 | 0/3 | Neg. |
| respiratory syncytial virus | Long A | Isolate | 0/3 | 0/3 | Neg. |
| rhinovirus |  | Isolate | 0/3 | 0/3 | Neg. |
| parainfluenza 1 | C35 | Isolate | 0/3 | 0/3 | Neg. |
| parainfluenza 2 | Greer | Isolate | 0/3 | 0/3 | Neg. |
| parainfluenza 3 | C-43 | Isolate | 0/3 | 0/3 | Neg. |
| parainfluenza 4 | M-25 | Isolate | 0/3 | 0/3 | Neg. |

45

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Endogenous Interference Substances Studies:

The CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel uses conventional well-established nucleic acid

extraction methods and based on our experience with CDC’s other EUA assays, including the CDC Novel

Coronavirus 2012 Real-time RT-PCR Assay for the presumptive detection of Middle East Respiratory

Syndrome Coronavirus (MERS-CoV) and the CDC Human Influenza Virus Real-Time RT-PCR Diagnostic  Panel-Influenza A/H7 (Eurasian Lineage) Assay for the presumptive detection of novel influenza A (H7N9)  virus that are both intended for use with a number of respiratory specimens, we do not anticipate  interference from common endogenous substances.

Specimen Stability and Fresh-frozen Testing:

To increase the likelihood of detecting infection, CDC recommends collection of lower respiratory and

upper respiratory specimens for testing. If possible, additional specimen types (e.g., stool, urine) should  be collected and should be stored initially until decision is made by CDC whether additional specimen  sources should be tested. Specimens should be collected as soon as possible once a PUI is identified  regardless of symptom onset. Maintain proper infection control when collecting specimens. Store

specimens at 2-8°C and ship overnight to CDC on ice pack. Label each specimen container with the  patient’s ID number (e.g., medical record number), unique specimen ID (e.g., laboratory requisition  number), specimen type (e.g., nasal swabs) and the date the sample was collected. Complete a CDC

Form 50.34 for each specimen submitted.

Clinical Performance:

As of February 22, 2020, CDC has tested 2071 respiratory specimens from persons under investigation

(PUI) in the U.S. using the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel. Specimen types include  bronchial fluid/wash, buccal swab, nasal wash/aspirate, nasopharyngeal swab, nasopharyngeal/throat  swab, oral swab, sputum, oropharyngeal (throat) swab, swab (unspecified), and throat swab.

46

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table 10: Summary of CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel  Data Generated by Testing Human Respiratory Specimens Collected from PUI Subjects in the U.S.

| Specimen Type | 2019 nCoV Negative | 2019 nCoV Positive | Inconclusive | Invalid | Total |
| --- | --- | --- | --- | --- | --- |
| Bronchial fluid/wash | 2 | 0 | 0 | 0 | 2 |
| Buccal swab | 5 | 1 | 0 | 0 | 6 |
| Nasal wash/aspirate | 6 | 0 | 0 | 0 | 6 |
| Nasopharyngeal swab | 927 | 23 | 0 | 0 | 950 |
| Nasopharyngeal swab/throat swab | 4 | 0 | 0 | 0 | 4 |
| Oral swab | 476 | 9 | 0 | 0 | 485 |
| Pharyngeal (throat) swab | 363 | 10 | 0 | 1 | 374 |
| Sputum | 165 | 5 | 0 | 0 | 170 |
| Swab (unspecified)1 | 71 | 1 | 0 | 0 | 72 |
| Tissue (lung) | 2 | 0 | 0 | 0 | 2 |
| Total | 2021 | 49 | 0 | 1 | 2071 |

1Actual swab type information was missing from these upper respiratory tract specimens.

Two thousand twenty-one (2021) respiratory specimens of the 2071 respiratory specimens tested  negative by the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel. Forty-nine (49) of the 2071

respiratory specimens tested positive by the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel. Only

one specimen (oropharyngeal (throat) swab) was invalid. Of the 49 respiratory specimens that tested

positive by the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel, seventeen (17) were confirmed by

genetic sequencing and/or virus culture (positive percent agreement = 17/17, 95% CI: 81.6%-100%)

During the early phase of the testing, a total of 117 respiratory specimens collected from 46 PUI subjects  were also tested with two analytically validated real-time RT-PCR assays that target separate and

independent regions of the nucleocapsid protein gene of the 2019-nCoV, N4 and N5 assays. The

nucleocapsid protein gene targets for the N4 and N5 assays are different and independent from the

nucleocapsid protein gene targets for the two RT-PCR assays included in the CDC 2019-nCoV Real-Time

RT-PCR Diagnostic Panel, N1 and N2. Any positive result from the N4 and/or the N5 assay was further  investigated by genetic sequencing.

Performance of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel testing these 117 respiratory  specimens was estimated against a composite comparator. A specimen was considered comparator  negative if both the N4 and the N5 assays were negative. A specimen was considered comparator

positive when the N4 and/or the N5 assay generated a positive result, and the comparator positive

result(s) were further investigated and confirmed to be 2019-nCoV RNA positive by genetic sequencing.

47

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table 11: Percent Agreement of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel with the  Composite Comparator

| CDC 2019-nCoV Panel Result | Composite Comparator Result |  |
| --- | --- | --- |
|  | Positive | Negative |
| Positive | 131 | 0 |
| Inconclusive | 0 | 0 |
| Negative | 0 | 104 |

1Composite comparator results were available for 13 of 49 CDC 2019-nCoV Panel positive specimens

only.

Positive percent agreement = 13/13 = 100% (95% CI: 77.2% - 100%)

Negative percent agreement = 104/104 = 100% (95% CI: 96.4% - 100%)

Enzyme Master Mix Evaluation:

The limit of detection equivalence between the Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix  and the following enzyme master mixes was evaluated: Quantabio qScript XLT One-Step RT-qPCR  ToughMix, Quantabio UltraPlex 1-Step ToughMix (4X), and Promega GoTaq® Probe 1- Step RT-qPCR  System. Serial dilutions of 2019 novel coronavirus (SARS CoV-2) transcript were tested in triplicate with

the CDC 2019-nCoV Real-time RT-PCR Diagnostic Panel using all four enzyme master mixes. Both

manufactured versions of oligonucleotide probe, BHQ and ZEN, were used in the comparison. The lowest

detectable concentration of transcript at which all replicates tested positive using the Quantabio qScript  XLT One-Step RT-qPCR ToughMix and Quantabio UltraPlex 1-Step ToughMix (4X) was similar to that

observed for the Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix. The lowest detectable  concentration of transcript when using the Promega GoTaq® Probe 1- Step RT-qPCR System was one

dilution above that observed for the other candidates when evaluated with the BHQ version of the CDC  assays. The candidate master mixes all performed equivalently or at one dilution below the Thermo

Fisher TaqPath™ 1-Step RT-qPCR Master Mix when evaluated with the ZEN version of the CDC assays.

Table 12: Limit of Detection Comparison for Enzyme Master Mixes – BHQ Probe Summary Results

| Copy Number | Thermo Fisher TaqPath™ 1-Step RT qPCR Master Mix |  | Quantabio qScript XLT One-Step RT-qPCR ToughMix |  | Quantabio UltraPlex 1- Step ToughMix (4X) |  | Promega GoTaq® Probe 1- Step RT-qPCR System |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 2019- nCoV_N1 | 2019- nCoV_N2 | 2019- nCoV_N1 | 2019- nCoV_N2 | 2019- nCoV_N1 | 2019- nCoV_N2 | 2019- nCoV_N1 | 2019- nCoV_N2 |
| 102 copies/µL | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 |
| 101 copies/µL | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 |
| 100 copies/µL | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 2/3 |
| 10-1 copies µL | 2/3 | 0/3 | 1/3 | 1/3 | 1/3 | 1/3 | 0/3 | 0/3 |

48

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table 13: Limit of Detection Comparison for Enzyme Master Mixes – ZEN Probe Summary Results

| Copy Number | Thermo Fisher TaqPath™ 1-Step RT qPCR Master Mix |  | Quantabio qScript XLT One-Step RT-qPCR ToughMix |  | Quantabio UltraPlex 1- Step ToughMix (4X) |  | Promega GoTaq® Probe 1- Step RT-qPCR System |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 2019- nCoV_N1 | 2019- nCoV_N2 | 2019- nCoV_N1 | 2019- nCoV_N2 | 2019- nCoV_N1 | 2019- nCoV_N2 | 2019- nCoV_N1 | 2019- nCoV_N2 |
| 102 copies/µL | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 |
| 101 copies/µL | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 |
| 100 copies/µL | 3/3 | 2/3 | 3/3 | 3/3 | 3/3 | 2/3 | 3/3 | 3/3 |
| 10-1 copies µL | 1/3 | 1/3 | 0/3 | 0/3 | 0/3 | 1/3 | 1/3 | 1/3 |

Retrospective positive (18) and negative (17) clinical respiratory specimens were extracted using the  QIAGEN EZ1 Advanced XL instrument and EZ1 DSP Virus Kit and were tested with the CDC 2019-nCoV  Real-time RT-PCR Diagnostic Panel using the Quantabio qScript XLT One-Step RT-qPCR ToughMix,

Quantabio UltraPlex 1-Step ToughMix (4X), and Promega GoTaq® Probe 1- Step RT-qPCR System master  mixes. All three enzyme master mixes performed equivalently, demonstrating 100% positive and 100%

negative agreement with expected results and a 95% confidence interval of 82.4%-100% and 81.6%- 100%, respectively.

Table 14: Clinical Comparison – Retrospective Study Summary Results

| CDC 2019-nCoV Real-time RT PCR Diagnostic Panel Result | Quantabio qScript XLT One-Step RT-qPCR ToughMix |  | Quantabio UltraPlex 1-Step ToughMix (4X) |  | Promega GoTaq® Probe 1- Step RT-qPCR System |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Positive | Negative | Positive | Negative | Positive | Negative |
| Positive | 18 | 0 | 18 | 0 | 18 | 0 |
| Negative | 0 | 17 | 0 | 17 | 0 | 17 |

Roche MagNA Pure 24 and MagNA Pure 96 Extraction Platform Evaluation:

Performance of the 2019-CoV Real-time RT-PCR Diagnostic Panel using the Roche MagNA Pure 24 and  MagNA Pure 96 extraction platforms was compared to performance with an authorized extraction  method. Serial dilutions of quantified inactivated SARS-CoV-2 virus (USA-WA1/2020; 100 RNA copies/µL)  in lysis buffer were added to pooled negative upper respiratory tract specimen matrix. Five samples of  each dilution were extracted in parallel with the QIAGEN EZ1 Advanced XL (EZ1 DSP Virus Kit Cat# 62724)

and the Roche MagNA Pure 24 (MagNA Pure 24 Total NA Isolation Kit Cat# 07658036001) and Roche

MagNA Pure 96 (MagNA Pure 96 DNA and Viral Nucleic Acid Small Volume Kit Cat# 06543588001)  extraction platforms and evaluated using the 2019-nCoV Real-Time RT-PCR Diagnostic Panel and  ThermoFisher TaqPath™ 1-Step RT-qPCR Master Mix. The observed LoD was defined as the lowest  concentration at which 100% (5 out of 5 total) of all replicates tested positive for both primer/probe sets

(N1 and N2) in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel. The acceptance criteria for

49

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

equivalence were defined as demonstrating an observed LoD either at the same endpoint or within a 3- fold dilution. The results showed that both the MagNA Pure 24 and MagNA Pure 96 extraction platforms

performed equivalently or within one 3-fold dilution of the LoD observed when using the QIAGEN EZ1

Advanced XL extraction platform.

Table 15. Limit of Detection Comparison between the QIAGEN EZ1 Advanced XL, Roche MagNA Pure  96, and Roche MagNA Pure 24 Extraction Platforms using the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel

| Platform | Parameter | 2019-nCoV_N1 Assay |  |  | 2019-nCoV_N2 Assay |  |  | Observed LoD1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QIAGEN EZ1 Advanced XL | RNA copies/µL | 101.0 | 100.5 | 100.0 | 101.0 | 100.5 | 100.0 | 100.5 |
|  | # pos./total | 5/5 | 5/5 | 5/5 | 5/5 | 5/5 | 3/5 |  |
|  | Mean Ct2 | 34.0 | 35.0 | 36.3 | 33.9 | 36.6 | NA |  |
|  | Std. Deviation | 0.2 | 0.8 | 0.2 | 0.4 | 0.9 | NA |  |
| Roche MagNA Pure 96 | RNA copies/µL | 101.0 | 100.5 | 100.0 | 101.0 | 100.5 | 100.0 | 100.5 |
|  | # pos./total | 5/5 | 5/5 | 5/5 | 5/5 | 5/5 | 2/5 |  |
|  | Mean Ct2 | 33.3 | 34.6 | 36.1 | 33.2 | 35.7 | NA |  |
|  | Std. Deviation | 0.5 | 0.5 | 0.3 | 0.3 | 0.4 | NA |  |
| Roche MagNA Pure 24 | RNA copies/µL | 101.0 | 100.5 | 100.0 | 101.0 | 100.5 | 100.0 | 101.0 |
|  | # pos./total | 5/5 | 3/5 | 3/5 | 5/5 | 5/5 | 5/5 |  |
|  | Mean Ct2 | 34.4 | NA | NA | 35.2 | 36.9 | 36.2 |  |
|  | Std. Deviation | 0.6 | NA | NA | 0.5 | 1.0 | 0.8 |  |

1Concentration is presented in RNA copies/µL. The observed LoD is the lowest concentration where both assays showed 100%  positive detection.

2Mean Ct reported for dilutions that show 100% positivity. Calculations only include positive results.  NA = not applicable

Previously characterized clinical remainder specimens (14 positive and 15 negative) were extracted using  both the Roche MagNA Pure 96 and MagNA Pure 24 extraction platforms and evaluated using the 2019-

nCoV Real-Time RT-PCR Diagnostic Panel and Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix.  Acceptance criteria for clinical equivalence was defined as demonstrating 100% concurrence with  qualitative results shown with the authorized comparator method (QIAGEN EZ1 Advanced XL). Results  from this study showed 100% concurrence with the comparator method for both the Roche MagNA Pure  96 and Roche MagNA Pure 24 extraction platforms when used with the CDC 2019-nCoV Real-Time RT PCR Diagnostic panel.

Table 16. Clinical Comparison Results – Retrospective Study Results

| Test Platform | Test Platform Result | QIAGEN EZ1 Advanced XL Result |  | Positive % Agreement (CI)1 | Negative % Agreement (CI)1 |
| --- | --- | --- | --- | --- | --- |
|  |  | Positive | Negative |  |  |
| Roche MagNA Pure 96 | Positive | 14 | 0 | 100.0 (78.5 – 100.0) | 100.0 (79.6 – 100.0) |
|  | Negative | 0 | 15 |  |  |
| Roche MagNA Pure 24 | Positive | 14 | 0 | 100.0 (78.5 – 100.0) | 100.0 (79.6 – 100.0) |
|  | Negative | 0 | 15 |  |  |

1 CI = 95% confidence interval

50

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Promega Maxwell® RSC 48 Extraction Platform Evaluation:

Performance of the 2019-CoV Real-time RT-PCR Diagnostic Panel using the Promega Maxwell® RSC 48

extraction platform was compared to performance with an authorized extraction method. Serial  dilutions of quantified inactivated SARS-CoV-2 virus (USA-WA1/2020; 100 RNA copies/µL) in VTM were  added to pooled negative upper respiratory tract specimen matrix. Five samples of each dilution were

extracted in parallel with the QIAGEN EZ1® Advanced XL (EZ1 DSP Virus Kit Cat# 62724) and the Promega

Maxwell® RSC 48 (Promega Maxwell® Viral Total Nucleic Acid Purification Kit Cat# AS1330) extraction

platforms and evaluated using the 2019-nCoV Real-Time RT-PCR Diagnostic Panel and ThermoFisher  TaqPath™ 1-Step RT-qPCR Master Mix. The observed LoD was defined as the lowest concentration at  which 100% (5 out of 5 total) of all replicates tested positive for both primer/probe sets (N1 and N2) in  the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel. The acceptance criteria for equivalence were

defined as demonstrating an observed LoD either at the same endpoint or within a 3-fold dilution. The  results showed that the performance of the Maxwell® RSC 48 extraction platform performed

equivalently or within one 3-fold dilution of the LoD observed when using the QIAGEN EZ1® Advanced XL  extraction platform.

Table 17. Limit of Detection Comparison Between the QIAGEN EZ1® Advanced XL and Promega  Maxwell® RSC 48 Extraction Platforms Using the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

| Platform | Parameter | 2019-nCoV_N1 Assay |  |  | 2019-nCoV_N2 Assay |  |  | Observed LoD1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| QIAGEN EZ1® Advanced XL | RNA copies/µL | 100.5 | 100.0 | 10-0.5 | 100.5 | 100.0 | 10-0.5 | 10 0.0 |
|  | # pos./total | 5/5 | 5/5 | 0/5 | 5/5 | 5/5 | 3/5 |  |
|  | Mean Ct2 | 32.27 | 33.80 | NA | 35.13 | 36.41 | NA |  |
|  | Std. Deviation | 0.81 | 0.40 | NA | 0.81 | 0.40 | NA |  |
| Promega Maxwell® RSC 48 | RNA copies/µL | 100.5 | 100.0 | 10-0.5 | 100.5 | 100.0 | 10-0.5 | 10 0.0 |
|  | # pos./total | 5/5 | 5/5 | 3/5 | 5/5 | 5/5 | 5/5 |  |
|  | Mean Ct2 | 31.11 | 32.97 | NA | 31.89 | 33.95 | 35.17 |  |
|  | Std. Deviation | 0.24 | 0.34 | NA | 0.24 | 0.35 | 0.65 |  |

1Concentration is presented in RNA copies/µL. The observed LoD is the lowest concentration where both assays showed 100%  positive detection.

2Mean cycle threshold (Ct) reported for dilutions that show 100% positivity. Calculations only include positive results.  NA = not applicable

Previously characterized clinical remainder specimens (15 positive and 15 negative) were extracted using  the Promega Maxwell® RSC 48 extraction platform alongside the currently authorized QIAGEN EZ1®

Advanced XL extraction platform and evaluated using the 2019-nCoV Real-Time RT-PCR Diagnostic Panel

and Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix. Results from the Maxwell® RSC 48 were  compared with the QIAGEN EZ1® Advanced XL extraction performed in parallel showing 100% (15/15)  qualitative concurrence on positive samples and 93.3% (14/15) qualitative concurrence on negative  samples. This evaluation showed that two originally negative (QIAGEN QIAamp® DSP Viral RNA Mini Kit)

specimens (Specimens 16 and 24) yielded an inconclusive result after extraction using the QIAGEN EZ1®

Advanced XL. Repeat of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel resolved one of the two  specimens (Specimen 24, negative result). The second specimen (Specimen 16) remained inconclusive.  Both these specimens yielded a negative result on the Maxwell® RSC 48.

51

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table 18. Clinical Comparison Results – Retrospective Study Results

| Test Platform |  | Promega Maxwell® RSC 48 |  |  | Positive % Agreement (CI)1 | Negative % Agreement (CI)1 |
| --- | --- | --- | --- | --- | --- | --- |
|  | Result | Positive | Negative | Inconclusive |  |  |
| QIAGEN EZ1® Advanced XL | Positive | 15 | 0 | 0 | 100.0 (79.6-100.0) | 93.3 (70.2-98.9) |
|  | Negative | 0 | 14 | 0 |  |  |
|  | Inconclusive | 0 | 1 | 0 |  |  |

1 CI = 95% confidence interval

Disposal

Dispose of hazardous or biologically contaminated materials according to the practices of your  institution.

References

1. Ballew, H. C., et al. “Basic Laboratory Methods in Virology,” DHHS, Public Health Service 1975  (Revised 1981), Centers for Disease Control and Prevention, Atlanta, Georgia 30333.

2. Clinical Laboratory Standards Institute (CLSI), “Collection, Transport, Preparation and Storage of  Specimens for Molecular Methods: Proposed Guideline,” MM13-A

3. Lieber, M., et al. "A Continuous Tumor Cell Line from a Human Lung Carcinoma with Properties of

Type II Alveolar Epithelial Cells." International Journal of Cancer 1976, 17(1), 62-70.

52

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Revision History

| Revision # | Effective Date | Summary of Revisions |
| --- | --- | --- |
| 1 | February 4, 2020 | Original Instructions for Use |
| 2 | March 15, 2020 | • Intended use update • Removal of N3 primer and probe set from Diagnostic Panel • Performance data update • Addition of alternative nucleic acid extraction platforms • Addition of acceptable alternatives to HSC and addition of QIAGEN RUO extraction reagents • Positive results no longer presumptive. No confirmation of positive results required |
| 3 | March 30, 2020 | • Addition of alternative enzyme master mix options |
| 4 | June 12, 2020 | • Addition of MagNA Pure 24 extraction method • Addition of performance data for the MagNA Pure 96 extraction method with SARS-CoV-2 • Addition of heat treatment alternative to specimen extraction • Addition of Roche and QIAGEN external lysis buffer alternatives • Acknowledgment of FDA policy permitting end users to qualify alternative components without seeking an EUA or EUA amendment |
| 5 | July 13, 2020 | • Addition of Promega Maxwell® RSC 48 extraction method • Update to in silico inclusivity analyses |
| 6 | December 1, 2020 | • Addition of specimen pooling instructions and monitoring procedure (Appendices B and C) • Addition of Promega Maxwell CSC 48 extraction instrument • Addition of data for CDC testing of FDA reference panel |

Contact Information, Ordering, and Product Support

For technical and product support, contact the CDC Division of Viral Diseases directly.  Send email to: respvirus@cdc.gov

Note: If your laboratory is using reagents sourced from someone other than the CDC International  Reagent Resource, please refer to the manufacturer’s instructions provided with the commercial  materials.

53

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Appendix A: Heat Treatment Alternative to Extraction

UltraPlex 1-Step ToughMix (4X)

This procedure is only for use by public health laboratories.

Purpose:

In response to a global shortage of nucleic acid extraction reagents causing significant delays in testing,

the CDC has investigated the use of a heat treatment method requiring minimal reagents as a specimen  processing alternative to nucleic acid extraction for use with the 2019-nCoV Real-Time RT-PCR Diagnostic  Panel.

Where possible, laboratories should use qualified RNA or total nucleic acid extraction methods for  processing of specimens for subsequent testing by the CDC 2019-nCoV Real-Time RT-PCR Diagnostic  Panel. Extraction removes inhibitory substances from specimens that could negatively impact PCR  performance.

This procedure for use of heat treatment for specimen processing is only recommended when a  shortage of qualified extraction reagents is a limiting factor in a laboratory’s ability to meet urgent

COVID-19 testing demand.

Precautions/Warnings/Limitations:

• CDC has evaluated this heat treatment process and has determined that this process is effective  for inactivation of SARS-CoV-2 in patient specimens.

• Performance was evaluated with only upper respiratory specimens. Heat treatment of lower

respiratory specimens for subsequent testing by the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel has not been evaluated.

• This procedure for heat treatment of specimens is only for use with the Quantabio UltraPlex 1- Step ToughMix (4X).

• Heat treatment should only be conducted when a lab is ready to test the specimens by PCR.  Testing of heat-treated specimens must be conducted the same day.

Acceptable Specimens:

• Upper respiratory specimens

Note: Do not use heat treatment to process specimens that appear bloody or that contain

particulate matter. Such specimens should be extracted using a qualified RNA or TNA extraction  method prior to testing.

Materials Required (not provided):

• 70% ethanol

• 10% bleach, freshly prepared

• 96-well PCR reaction plates (Applied Biosystems catalog # 4346906, 4366932, 4346907, or

equivalent)

• Optical strip caps (Applied Biosystems 4323032, or equivalent)

54

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

• 1.5 mL Sarstedt tubes or equivalent

• Aerosol resistant micropipette tips

• Micropipettes

• 96-well cold block

• Cold blocks for 1.5 mL - 2.0 mL tubes

• Vortex mixer

• 96-well plate centrifuge or equivalent

• Thermal cycler or equivalent

• Class II Biological Safety Cabinet (BSC)

Procedure:

Sample Preparation

1) Decontaminate BSC with 10% bleach followed by 70% ethanol.

2) If samples are frozen, thaw on ice or at 4°C. Wipe the outside of the sample tube with 70%

ethanol. Place thawed sample on cold rack or ice in BSC.

3) Pulse vortex each sample and briefly spin down in a centrifuge to collect the liquid at the  bottom of the tube.

Heat Treatment

1) Place a thermal cycler in the BSC, turn on, and program for 95°C for 1 min followed by 4°C  hold.

2) Place a 96-well PCR plate onto a cold rack or ice in the BSC.

3) Transfer 100 µL of each sample to the 96-well PCR plate and securely cap each well using

optical strip caps.

NOTE: Ensure that an HSC extraction control is included in each batch run as required under  CLIA.

4) Place this 96-well PCR plate on the pre-heated thermal cycler and start run. Leave plate on

thermal cycler at 4°C, or place on ice or a cold block.

5) Remove plate and centrifuge for 1 minute at 500 x g to pellet cellular debris.

6) Place plate on a cold rack or ice and proceed to testing the supernatant by rRT-PCR.  7) Testing of heat-treated specimens must be conducted the same day heat treatment is  performed. For long term storage, keep the original specimen at ≤-70°C.

Special Testing Considerations for Heat Treated Specimens:

Enzyme Master Mix

• Testing of specimens that have been processed with heat treatment should be conducted  with the Quantabio UltraPlex 1-Step ToughMix (4X), which demonstrated the best  performance with heat treated specimens. PCR testing of heat-treated specimens should  follow the instructions in the main body of this Instructions for Use document.

55

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Resolution of Inconclusive and Invalid Results

• Retesting of heat-treated specimens that generated an inconclusive or invalid result must  include extraction of the original specimen with a qualified RNA or total nucleic acid (TNA)  extraction method, if available. Do not re-test the heat-treated specimen material to resolve

inconclusive or invalid test results.

Verification:

CDC recommends performance of verification studies for the heat treatment method prior to diagnostic

use that includes side-by-side preparation of a panel of positive and negative clinical specimens using a

qualified extraction method and this heat treatment method with subsequent testing by the CDC 2019- nCoV Real-Time RT-PCR Diagnostic Panel.

Performance Characteristics:

Quantabio UltraPlex 1-Step ToughMix (4X)

Limit of Detection Comparison

Serial dilutions of inactivated SARS-CoV-2 [SARS-CoV-2 USA-WA1/2020] were prepared in simulated

specimen material (human A549 cells suspended in viral transport medium). Each concentration was

prepared side-by-side five times by both EZ1 extraction and by heat treatment. Each extracted or heat treated sample was subsequently tested by the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel using  the Quantabio UltraPlex 1-Step ToughMix (4X) on the Applied Biosystems 7500 Fast Dx instrument.

Observed detection was similar between the two specimen preparation methods.

Table 1: UltraPlex Limit of Detection Comparison between QIAGEN EZ1 Advanced XL extraction and heat  treatment (95°C for 1 min) method – Summary Results

| Enzyme | Platform | Parameter | 2019-nCoV_N1 Assay |  |  |  |  | 2019-nCoV_N2 Assay |  |  |  |  | Observed LoD1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| - n 1 ) o i x X t i e 4 l d ( P d x a i A r t M e l t h U a g l o u p i o b m T a e t p T n e a L t u S µ Q 5 | d N e c E 1 L n G Z a X A E v I d Q A | RNA copies/µL | 101.0 | 100.5 | 100.0 | 10-0.5 | 10-1.0 | 101.0 | 100.5 | 100.0 | 10-0.5 | 10-1.0 | 100.5 |
|  |  | # pos./total | 5/5 | 5/5 | 4/5 | 4/5 | 3/5 | 5/5 | 5/5 | 5/5 | 2/5 | 2/5 |  |
|  |  | Mean Ct2 | 34.11 | 34.59 | NA | NA | NA |  | 32.97 33.76 34.70 |  | NA | NA |  |
|  |  | Std. Deviation | 0.75 | 0.99 | NA | NA | NA | 0.33 | 0.72 | 0.98 | NA | NA |  |
|  | t n n i e t C m a ° m t e 5 1 a 9 H r e o r f T | RNA copies/µL | 101.0 | 100.5 | 100.0 | 10-0.5 | 10-1.0 | 101.0 | 100.5 | 100.0 | 10-0.5 | 10-1.0 | 100.5 |
|  |  | # pos./total | 5/5 | 5/5 | 4/5 | 5/5 | 1/5 | 5/5 | 5/5 | 4/5 | 2/5 | 1/5 |  |
|  |  | Mean Ct2 | 33.41 | 34.32 | NA | 36.73 | NA |  | 33.45 35.25 | NA | NA | NA |  |
|  |  | Std. Deviation | 0.62 | 0.40 | NA | 0.82 | NA | 0.40 | 0.80 | NA | NA | NA |  |

1Concentration is presented in RNA copies/µL. The observed LoD is the lowest concentration where both assays showed 100%  positive detection.

2Mean Ct reported for dilutions that show 100% positivity. Calculations only include positive results.

NA = not applicable

56

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Clinical Comparison

A panel of 39 upper respiratory specimens were tested side-by-side using extraction with the Qiagen EZ1

extraction instrument and heat treatment. Extracted and heat-treated specimens were subsequently  tested with the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel using the Quantabio UltraPlex 1-Step

ToughMix (4X). Qualitative results were compared to demonstrate agreement.

Table 2: Clinical Comparison Results Summary – Heat Treatment versus QIAGEN EZ1 Advanced XL

|  | Test Result | Heat Treatment |  |  | Total | Positive % Agreement (CI)1 | Negative % Agreement (CI)1 |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Positive | Inconclusive | Negative |  |  |  |
| QIAGEN EZ1 Advanced XL | Positive | 18 | 1 | 0 | 19 | 94.7 (75.4-99.1) | 100 (83.9-100) |
|  | Inconclusive | 0 | 0 | 0 | 0 |  |  |
|  | Negative | 0 | 0 | 20 | 20 |  |  |
|  | Total | 18 | 1 | 20 | 39 |  |  |

1 CI = 95% confidence interval

Questions and Comments:

If you have questions or comments about this procedure, please send by email to: respvirus@cdc.gov

57

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Appendix B: Pooled Specimen Preparation and Processing

Purpose:

In response to strong demand for higher throughput testing approaches as well as a global shortage of  nucleic acid extraction reagents causing significant delays in testing, the CDC has evaluated specimen  pooling and determined that pooling of up to 4 specimens is suitable for use with the 2019-nCoV Real Time RT-PCR Diagnostic Panel.

Specimen pooling may cause a slight reduction in test sensitivity and therefore may be most appropriate  for screening or diagnostic testing when laboratory staff, equipment or reagents are insufficient to  accommodate testing demand. Specimen pooling only presents a throughput advantage when the

disease prevalence is low. Therefore, laboratories should monitor specimen positivity rates over time to  determine if pooling of specimens continues to provide a test throughput advantage over individual

specimen testing. Please see Appendix C for more detail on implementation considerations.

While this procedure describes the process to prepare, process and test a pool size of up to 4 specimens,

specimen pool sizes from 2-4 are authorized for use with the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel. When using a pool size less than 4 specimens, please use the following instructions as

a model. Pooled specimen input volume and the pooled specimen to lysis buffer volume ratios must

remain as prescribed below (not a lower proportion of lysis buffer) to ensure inactivation of SARS-CoV-2  in patient specimens. An N-pool specimen approach should include equal volumes of each of the N

specimens pooled together to create the total pooled specimen input volume required under the below

pooled specimen extraction instructions.

Precautions/Warnings/Limitations:

• Pooling of specimens has the potential to decrease sensitivity. The specimens in a pooled  procedure are diluted, which could result in a low concentration of viral genetic material below

the limit of detection of a given test.

• When pooling specimens, the laboratory cannot ensure the diagnostic integrity of an individual  specimen because it is combined with other specimens before testing. Specimen integrity can be

affected by the quality of swab specimen collection, which could result in some swabs having

limited amounts of viral genetic material for detection. Inadequate individual specimens,  including those with limited amounts of viral genetic material, might not be eliminated from the

pooled specimen before testing and may be reported as negative through this process.

• Performance of this specimen pooling process was evaluated with nasopharyngeal (NP) swabs.

Pooling of other types of specimens for subsequent testing with the CDC 2019-nCoV Real-Time  RT-PCR Diagnostic Panel has not been evaluated, however evidence suggests the Ct value

distribution for other upper respiratory swab specimens is similar to that observed for NP swab  specimens.

• Pooling should only be implemented in laboratories with a minimum of 10 days experience using

the CDC 2019-nCoV Real-time RT-PCR Diagnostic Panel for diagnostic testing in the population of  specimens being considered for pooling.

58

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

• This procedure for pooling of specimens has only been evaluated for use with the QIAGEN EZ1

Advanced XL, Roche MagNA Pure 96, and the Promega Maxwell RSC 48 instruments. Pooling

using other extraction methods is not authorized under this EUA.

• Sufficient volume of specimens must be available to allow subsequent extraction and testing of

individual specimens should the specimen pool test positive, inconclusive or invalid.  • Appropriate interpretation, reporting and next steps may be different for surveillance testing and

are not limited by the diagnostic interpretation instructions below.

• Laboratories are encouraged to revisit the question of whether pooling continues to make sense  when testing demand no longer exceeds laboratory capacity and/or when testing reagents are no  longer in short supply.

Acceptable Specimen Types:

• Upper respiratory swabs (e.g., nasopharyngeal (NP) swabs, oropharyngeal (OP) swabs, NP/OP

swabs combined, nasal swabs)

Note: Specimens must have adequate volume to support pooled specimen testing and any  subsequent individual re-testing.

59

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Overview of Pooled Specimen Testing:

This flow chart provides an overview of the steps involved in pooled specimen testing and location of the  instructions for each step within the document. For those steps that are identical for both individual

specimen testing and for pooled specimen testing, instructions are found in the main body of the

instructions for use. For steps in which pooled specimen testing differs from individual specimen testing,

instructions specific to pooled specimen testing are presented in this Appendix.

60

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Procedure:

Specimen Pool Preparation

1) Decontaminate BSC with 10% bleach followed by 70% ethanol.

2) If specimens are frozen, thaw on ice or at 4°C. Wipe the outside of the specimen tube with 70%

ethanol. Place thawed specimen on cold rack or ice in BSC.

3) Pulse vortex each specimen and briefly spin down in a centrifuge to collect the liquid at the

bottom of the tube.

4) For each specimen pool, add an equal volume of each specimen (depending on total volume

needed for extraction platform) to be included into the pool, into a sterile, nuclease-free tube.

5) Pulse vortex each specimen pool and briefly centrifuge to collect the liquid at the bottom of the  tube.

6) Proceed to extraction using the modified extraction parameters below.

Extraction Instructions

NOTE: These pooled specimen extraction instructions have been modified to optimize

recovery of nucleic acid when processing pooled specimens. When testing individual

specimens, please follow the individual specimen extraction instructions (p. 14-16).

QIAGEN EZ1 Advanced XL

• Kit: Qiagen EZ1 DSP Virus Kit and Buffer AVL (supplied separately) for offboard lysis

• Card: EZ1 Advanced XL DSP Virus Card

• Instructions: Add 200 μL of pooled specimen to 200 μL of pre-aliquoted Buffer AVL (total

input sample volume is 400 μL). Proceed with the extraction on the EZ1 Advanced XL. Elution

volume is 60 μL.

Roche MagNA Pure 96

• Kit: Roche MagNA Pure 96 DNA and Viral NA Small Volume Kit

• Protocol: Viral NA Plasma Ext LysExt Lys SV 4.0 Protocol or Viral NA Plasma Ext Lys SV Protocol

• Instructions: Add 225 μL of pooled specimen to 225 μL of pre-aliquoted MagNA Pure 96  External Lysis Buffer (supplied separately) for total input sample volume of 450 μL. Proceed  with the extraction on the MagNA Pure 96. (Internal Control = none).

• Elution volume is 50 μL.

Promega Maxwell RSC 48 or Maxwell CSC 48

• Kit: Promega Maxwell® Viral Total Nucleic Acid Purification Kit

• Protocol: Viral Total Nucleic Acid

• Instructions: Add 240 μL of pooled specimen to 660 μL of pre-aliquoted External Lysis Buffer

(600 μL Lysis Buffer plus 60 μL Proteinase K; supplied within the kit) (total input volume is 900

μL). Proceed with the extraction on the Maxwell® RSC 48 or Maxwell CSC 48.

• Elution volume is 50 μL.

61

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Testing, Interpretation, and Reporting Instructions

1) After extraction, continue PCR testing using the steps outlined in the following sections of the

Instructions for Use: Assay Set Up (p. 16-20), Running a Test (p. 29-32) and Data Analysis

(p. 33-34).

2) Interpret controls and PCR test results as outlined in the following sections of the Instructions  for Use: Interpretation of Results and Reporting (p. 35-37), 2019-nCoV rRT-PCR Diagnostic  Panel Results Interpretation Guide (p. 38). All extraction and PCR run controls must perform

as expected.

A. If a specimen pool generates negative results, report “Not Detected” for each specimen  included in the pool. The test result report should include a statement that the specimen

was tested in a pooled format. No further testing is required.

B. If a specimen pool generates a positive result, do not report the result. Test each  specimen included in the pool individually, according to the Instructions for Use. Report  the individual result for each specimen.

C. If a specimen pool generates an inconclusive result, do not report the result. Test each

specimen included in the pool individually, according to the Instructions for Use. Report  the individual result for each specimen.

D. If a specimen pool generates an invalid result, do not report the result. Test each

specimen within the pool individually, according to the Instructions for Use. Report the  individual result for each specimen.

Verification:

Verification should be conducted by all laboratories prior to implementing pooling for diagnostic testing.

In addition to method verification studies, verification must include a validation of the accuracy and  function of the laboratory’s process of specimen and pool labeling, tracking and reporting through the  laboratory information management system(s).

Performance Characteristics:

Pooling Validation

A panel of 20 positive, previously characterized nasopharyngeal (NP) specimens, with 25% (5)

representing the weak positive range (Ct 36.00-39.99), were used for evaluation of specimen pooling  with the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel; 80 negative NP specimens were also

included. These specimens were used to create twenty positive 4-specimen pools, by combining one

positive specimen with three negative specimens. Twenty negative pools were also created by

combining equal volumes of four negative specimens. These pools were tested and qualitative results

from pooled-testing were compared with the results from individual-testing to determine the positive  and negative percent agreement of the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel for testing  using 4-specimen pools. The performance of 4-specimen pooling was evaluated using the QIAGEN EZ1®

Advanced XL, the Promega Maxwell RSC 48, or the Roche MagNA Pure 96 platforms. Modifications were

made to the input and output volume for each extraction platform, for pooling patient specimen, to

compensate for the dilution effect due to pooling.

62

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

QIAGEN EZ1® Advanced XL

Individual positive panel members and positive and negative pools were extracted using the QIAGEN

EZ1® Advanced XL platform for subsequent testing on the Applied Biosystems 7500 Fast  Dx with the Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix, CG. Results obtained for twenty, 4-

specimen pools were compared with the results from individual testing. Pools containing a positive

specimen were considered to be in agreement with results from individual testing if the pool

generated positive or inconclusive results (i.e., at least one of the SARS-CoV-2-specific primer and probe  sets generated a Ct value less than 40). Negative specimen pools were considered to be in agreement

with expected results if the pooled specimen result was negative (i.e., neither SARS-CoV-2-specific  primer and probe set generated positive results).

A comparison of qualitative test results from individual testing of panel members with pooled specimen

testing is summarized below in Table 1, along with testing agreement with results from individual testing  for positive and negative specimen pools. Of the five positive specimen panel members in the weak

positive category (determined by original diagnostic results): one returned an inconclusive result when  tested individually, but returned a positive result when tested in a 4-specimen pool; one returned a

negative result when tested individually, but returned a positive result when tested in a 4-specimen  pool; and three were in agreement when tested individually and in 4-specimen pools.

Table 1 -Summary of Pooled vs. Individual-testing - QIAGEN EZ1® Advanced XL

| Test Platform |  | Individual testing result |  |  | Positive Pools Percent Agreement (CI)1 | Negative Pools Percent Agreement (CI)1 |
| --- | --- | --- | --- | --- | --- | --- |
|  | Result | Positive | Inconclusive† | Negative |  |  |
| 4-specimen Pooling Result | Positive | 17 | 1 | 1 | 100% (82.4-100) | 95.2% (77.3-99.2) |
|  | Inconclusive‡ | 1 | 0 | 0 |  |  |
|  | Negative | 0 | 0 | 20 |  |  |

1CI = 95% confidence interval

†Inconclusive individually-tested results are not included in the final performance calculations.

‡Inconclusive pooled results are considered in agreement with the positive individually-tested results for the final performance calculations.

63

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Promega Maxwell RSC 48

Individual positive panel members and positive and negative pools were extracted using the Promega

Maxwell RSC 48 platform for subsequent testing on the Applied Biosystems 7500 Fast  Dx with the Thermo Fisher TaqPath™ 1-Step RT-qPCR Master Mix, CG. Results obtained for twenty, 4-

specimen pools were compared with the results from individual testing. Pools containing a positive

specimen were considered to be in agreement with the results from individual testing if the pooled  specimen generated positive or inconclusive results (i.e., at least one of the SARS-CoV-2-specific primer  and probe sets generated a Ct value less than 40). Twenty negative specimen pools (made from 4

known negative specimens) were considered to be in agreement with expected results if the pooled  specimen result was negative (i.e., neither SARS-CoV-2-specific primer and probe set generated positive

results).

A comparison of qualitative test results from individual testing of panel members with pooled specimen

testing is summarized below in Table 2, along with testing agreement with results from individual testing

for positive and negative specimen pools. Of the five positive specimen panel members in the weak  positive category (determined by original diagnostic results): one returned a positive result when tested  individually, but returned a negative result when tested in a 4-specimen pool; and four were in

agreement when tested individually and in 4-specimen pools. Additionally, one moderate to high

positive panel member returned an inconclusive result when tested individually but returned a positive  result when tested in a 4-specimen pool.

Table 2 -Summary of Pooled vs. Individual-testing Results – Promega Maxwell RSC 48

| Test Platform |  | Individual-testing result |  |  | Positive Pools Percent Agreement (CI)1 | Negative Pools Percent Agreement (CI)1 |
| --- | --- | --- | --- | --- | --- | --- |
|  | Result | Positive | Inconclusive† | Negative |  |  |
| 4-specimen Pooling Result | Positive | 16 | 1 | 0 | 94.7% (75.4-99.1) | 100% (83.9-100) |
|  | Inconclusive‡ | 2 | 0 | 0 |  |  |
|  | Negative | 1 | 0 | 20 |  |  |

1CI = 95% confidence interval

†Inconclusive individually-tested results are not included in the final performance calculations.

‡Inconclusive pooled results are considered in agreement with the positive individually-tested results for the final performance

calculations.

MagNA Pure 96

Individual positive panel members and positive and negative pools were extracted on the MagNA Pure

96 platform for subsequent testing on the Applied Biosystems 7500 Fast Dx using the Thermo

Fisher TaqPath™ 1-Step RT-qPCR Master Mix, CG. Results obtained for twenty, 4-specimen pools were  compared with the results from individual testing. Pools containing a positive specimen were

considered to be in agreement with the results from individual testing if the pooled specimen generated

positive or inconclusive results (i.e., at least one of the SARS-CoV-2-specific primer and probe sets

generated a Ct value less than 40). Twenty negative specimen pools (made from 4 known negative

specimens) were considered to be in agreement with expected results if the pooled specimen result was

negative (i.e., neither SARS-CoV-2-specific primer and probe set generated positive results).

64

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

A comparison of qualitative test results from individual testing of panel members with pooled specimen

testing is summarized below in Table 3, along with testing agreement with results from individual testing

for positive and negative specimen pools. Of the five positive specimen panel members in the weak  positive category (determined by original diagnostic results): one returned an inconclusive result when

tested individually, but returned a positive result when tested in a 4-specimen pool; and four were in

agreement when tested individually and in 4-specimen pools.

Table 3 -Summary of Pooled vs. Individual-testing Results – Roche MagNA Pure 96

| Test Platform |  | Individual-testing result |  |  | Positive Pools Percent Agreement (CI)1 | Negative Pools Percent Agreement (CI)1 |
| --- | --- | --- | --- | --- | --- | --- |
|  | Result | Positive | Inconclusive† | Negative |  |  |
| 4-specimen Pooling Result | Positive | 18 | 1 | 0 | 100% (83.2-100) | 100% (83.9-100) |
|  | Inconclusive‡ | 1 | 0 | 0 |  |  |
|  | Negative | 0 | 0 | 20 |  |  |

1CI = 95% confidence interval

†Inconclusive individually-tested results are not included in the final performance calculations.

‡Inconclusive pooled results are considered in agreement with the positive individually-tested results for the final performance

calculations.

In Silico Sensitivity

An in silico analysis was conducted to evaluate the effect of 4-sample pooling on the clinical sensitivity of  the CDC 2019-Novel Coronavirus Real-Time RT-PCR Diagnostic Panel using the QIAGEN EZ1® Advanced

XL extraction platform. This analysis was conducted by performing a Passing-Bablok regression using the  “Pooling Validation” data to calculate the Ct shifts resulting from the dilution effect of 4-specimen pools  (1 positive specimen combined with 3 negative specimens) for each target. In the regression analysis,

the X-axis displayed individual Ct values for positive specimens and the Y-axis displayed Ct values for the

corresponding pools with one positive specimen and 3 negative specimens. The regression analysis was  used to calculate an interval of Ct values [X*, 40] where individual specimens with Ct values within this

interval would have negative results in 4-specimen pools (1 positive and 3 negative).

The results from individually-tested NP swab specimens processed at the CDC using the QIAGEN EZ1®  Advanced XL extraction platform (n= 381), were analyzed to determine an in silico PPA for 4-specimen

pooling. The number of individual specimens with Ct values ranging from [X*, 40] was determined. The

X* value for the N1 target was 37.2 and the X* value for the N2 target was 37.5. The results, summarized

Table 4, show that 97.1% (370/381 95% CI 94.9-98.3%) of the specimens would not have negative results  when combined into 4-specimen pools.

65

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Table 4. in silico Sensitivity for Pooled Specimens Extracted using the QIAGEN EZ1® Advanced XL

| n | N1 Interval [X*, 40] | Number of samples with N1 Ct values in the interval | N2 Interval [X*, 40] | Number of samples with N2 Ct values in the interval | Number of Samples with both target Ct values in intervals | Neg | Inc | Pos | % Positive Percent Agreement* |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 381 | [37.2, 40] | 16 | [37.5, 40] | 63 | 11 | 11 | 57 | 313 | 97.1 |

*Since any pool that is not negative is re-tested as individual samples, the Positive Percent Agreement includes all pools that were not

negative.

The specimens submitted to CDC for testing included many that had generated inconclusive results in

the hands of public health laboratories and were forwarded to CDC for further testing. Thus, specimens  received by CDC for testing overrepresent the weak positive category. The CDC also obtained the results

of individually-tested NP swab specimens from three additional geographic locations: state lab 1 (n=  2217), state lab 2 (n= 6559), and state lab 3 (n= 2315). These data sets were analyzed with regard to the  percent of weak positives (Ct 36.00-39.99) collected at each site. The results, summarized in Table 5,

suggest that the percent of weak positive NP specimens received at State Health Department

laboratories is smaller than that obtained at CDC. Therefore, the in silico analysis, summarized in Table 4,

presents a PPA in a scenario where 4-specimen pooling is applied to testing populations with a greater  than average number of weak positive specimens.

Table 5. Percent of Weak Positive Specimens at CDC Compared with Three Geographic Locations

|  | n | Ct < 36.0 |  | Ct 36.0 - 39.99 |  |
| --- | --- | --- | --- | --- | --- |
|  |  | N1 | N2 | N1 | N2 |
| CDC | 381 | 314 (82.4 %) | 254 (66.7 %) | 67 (17.6 %) | 127 (33.3 %) |
| State 1 | 2217 | 2121 (95.7 %) | 2021 (91.2 %) | 150 (6.8 %) | 250 (11.3%) |
| State 2 | 6559 | 6343 (96.7 %) | 5883 (89.7 %) | 216 (3.3%) | 676 (10.3%) |
| State 3 | 2315 | 2073 (89.5 %) | 1884 (81.4 %) | 242 (10.5 %) | 431 (18.6 %) |

References:

1. Abdalhamid, B., et al. “Assessment of Specimen Pooling to Conserve SARS CoV-2 Testing

Resources.” Am J Clin Path 2020, 153, 715-718.

Questions and Comments:

If you have questions or comments about this procedure, please send by email to: respvirus@cdc.gov

66

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Appendix C: Implementation and Monitoring of Pooled Specimen Testing

Purpose:

These instructions are intended to assist laboratories and jurisdictions to implement and monitor the  efficiency and effectiveness of pooled specimen testing over time. A laboratory should have a minimum

of 10 days experience performing the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel for diagnostic

testing within the intended population prior to implementation of a pooling strategy.  Laboratories should weigh the potential risks associated with pooling as described in Appendix B against

the benefits of pooling for their laboratory and patient population. When pooling has been implemented

to address overwhelming testing demands and/or scarcity of testing materials, laboratories should  consider a return to individual specimen testing when laboratory capacity and resources permit.

Prior to Implementation of a Specimen Pooling Strategy:

• Determine the individual specimen positivity rate (PIndividual) for your testing population.  Before implementation of specimen pooling, evaluate test data from the testing population for

the previous 7-10 days to estimate the positivity rate (PIndividual) which is the number of positive

results divided by the total number of specimens tested during these 7-10 days. PIndividual will be

used to determine if pooling should be considered and will be used to monitor the performance

of your pooling strategy.

Note: To calculate the efficiency of 4-sample pooling, using PIndividual, apply the formula F=1/(1+  1/n-(1- PIndividual)n), where F is the efficiency and n is the pool size. For example, when PIndividual is

1%, the efficiency, F, is 3.46 for n = 4. This means that 1,000 tests can cover testing of 3,460  patients on average, which translates to a reduction in test volume of 71%. Using this formula, a

pool size of n = 3 produces a greater efficiency than n = 4 at PIndividual ≥ 13%.

• Identify your positivity thresholds.

Table 1: Association Between Individual Specimen Positivity Rates and Reduction in Test Volume with Pooling

| Specimen Positivity Rate (P) | No. tests with 4 specimen pooling for 300 patients | Reduction in testing volume with 4 specimen pooling | No. tests with 3 specimen pooling for 300 patients | Reduction in testing volume with 3 specimen pooling | No. tests with 2 specimen pooling for 300 patients | Reduction in testing volume with 2 specimen pooling |
| --- | --- | --- | --- | --- | --- | --- |
| 1% | 87 | 71% | 109 | 64% | 156 | 48% |
| 2% | 98 | 67% | 118 | 61% | 162 | 46% |
| 3% | 109 | 64% | 126 | 58% | 168 | 44% |
| 4% | 120 | 60% | 135 | 55% | 174 | 42% |
| 5% | 131 | 56% | 143 | 52% | 179 | 40% |
| 6% | 141 | 53% | 151 | 50% | 185 | 38% |
| 8% | 160 | 47% | 166 | 45% | 191 | 36% |
| 10% | 178 | 41% | 181 | 40% | 207 | 31% |
| 11% | 187 | 38% | 189 | 37% | 212 | 29% |
| 12% | 195 | 35% | 196 | 35% | 218 | 27% |
| 13% | 203 | 32% | 202 | 33% | 223 | 26% |

67

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

| Specimen Positivity Rate (P) | No. tests with 4 specimen pooling for 300 patients | Reduction in testing volume with 4 specimen pooling | No. tests with 3 specimen pooling for 300 patients | Reduction in testing volume with 3 specimen pooling | No. tests with 2 specimen pooling for 300 patients | Reduction in testing volume with 2 specimen pooling |
| --- | --- | --- | --- | --- | --- | --- |
| 14% | 211 | 30% | 209 | 30% | 228 | 24% |
| 15% | 218 | 27% | 216 | 28% | 233 | 22% |
| 16% | 226 | 25% | 222 | 26% | 238 | 21% |
| 17% | 233 | 22% | 228 | 24% | 243 | 19% |
| 18% | 239 | 20% | 235 | 22% | 248 | 17% |
| 19% | 246 | 18% | 241 | 20% | 253 | 16% |
| 20% | 252 | 16% | 246 | 18% | 258 | 14% |

Note: Table 1 does not account for running the positive, negative, and extraction controls with each plate.  Note: Unshaded cells indicate the optimal reduction in testing volume compared to other the pool sizes.

Table 1 presents the estimated increase in efficiency achieved through pooling at different rates

of specimen positivity. Laboratories are encouraged to use this table and factor in additional  testing necessary to monitor the effectiveness of pooling to determine the specimen positivity

thresholds (PThreshold) at which (1) four-specimen pooling is appropriate, (2) smaller specimen

pools are appropriate, and (3) the threshold at which specimen pooling offers no testing  capacity/reagent consumption advantages. These thresholds should be documented for your

laboratory’s pooling monitoring process.

Example: a laboratory might consider a 30% reduction in test volume as the minimum required

reduction in test volume to be considered beneficial (this is subjective). Thus, the laboratory  would set a positivity threshold (PThreshold) at 12%2 for 4 specimen pooling, 14% for 3 specimen

pooling, and 10% or less for 2 specimen pooling.

NOTE: It is recommended that PThreshold should not be greater than 20% for any pool size.

Moreover, laboratories may choose a threshold positivity rate lower than 20% for when to  consider a pooling strategy.

• Define and document a specimen pooling strategy/process for your laboratory.

• Verify laboratory performance of pooled specimen testing (CLIA requirement).

• Establish and implement a pooling monitoring process appropriate to assure the quality of results

generated through the laboratory’s pooling strategy. The pooling monitoring process should

document the laboratory’s PThreshold for each pool size, PIndividual and the laboratory’s process to

address the monitoring requirements below.

• Please note that information from the monitoring process described in this Appendix may not be the  only information useful to monitor for pooling performance issues or to aid in decision-making for  pooling strategy. Thus, laboratories are encouraged to monitor other sources of COVID-19  prevalence or test positivity rate information for their target population, comparing those trends to

2At PIndividual ≥ 13%, n=3 offers a greater efficiency than n=4, therefore it is recommended that PThreshold should not be greater  than 12% for n=4.

68

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

their pooling testing trends, as part of the laboratory’s pooling monitoring process in addition to the

approach below.

Monitoring of the Pooling Strategy

At regular intervals (weekly at a minimum, continual analysis with a moving 7-10 day average  preferred), laboratories must monitor the positivity rate of specimen pools as described below.

Figure 1: Overview of Pooling Monitoring Process to be Conducted at Defined Intervals

Assess current testing demand and reagent availability.

Will laboratory continue pooled specimen testing?

Yes No

Calculate PPool for the interval. Stop pooling.

No

No

Is PPool lower than PThreshold for the number  of specimens being pooled?

Yes

Is PPool less than 0.85 x PIndividual?

Reduce pool size  or stop pooling.

Yes

Conduct re-assessment of pooling.  Did re-assessment indicate pooling is still  appropriate?

Yes

Set new PIndividual based on re-assessment  data. Document for future monitoring  activities.

Continue pooled

specimen testing.

69

No

Stop pooling.

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Data Monitoring Steps to Be Conducted at Each Interval

1. Evaluate need for pooled specimen testing going forward.

Evaluate test demand and testing resource availability. When testing capacity and resource  availability are sufficient to meet testing demand without the use of specimen pooling, consider  whether the risks of reduced test sensitivity with pooling and extra effort of pooling outweigh the  benefits of pooled specimen testing.

2. Calculate Pooled Percent Positive (PPool) for the interval.

PPool = Specimens tested under pooling strategy generating positive results

Total specimens tested under pooling strategy

3. Check Interval PPool Against Positivity Threshold (PThreshold) for the pool size in use.  Compare the PPool for the current interval to the specimen positivity threshold established during

pooling implementation.

• If PPool is greater than PThreshold for the pool size in use, consider a reduction in pool size or

discontinuation of specimen pooling until specimen positivity rates decrease in your  testing population.

• If PPool is less than PThreshold, proceed to step 4.

4. Check Interval PPool Against Re-assessment Criteria

Compare PPool to PIndividual.

• If PPool is equal to or greater than 85% of PIndividual. (PPools ≥ 0.85 × PIndividual), then

monitoring checks are complete for this interval. Proceed with the current pooling

strategy.

• If PPool is less than 85% of PIndividual. (PPools < 0.85 × PIndividual), re-assessment is necessary to  determine if pooled-specimen testing process is still acceptable. Please proceed to the

Re-assessment of Pooling section below.

5. Update PIndividual if Indicated by Re-assessment.

If re-assessment is conducted and indicates n-specimen pooling is still acceptable, re-establish  PIndividual in your laboratory using the data generated during re-assessment.

If 10 positive specimens were tested during re-assessment:

10 positive specimens used for re-assessment 10 PIndividual = X The total number of specimens tested to reach 10 positive specimens 11  If 20 positive specimens were tested during re-assessment:

20 positive specimens used for re-assessment 20 PIndividual = X The total number of specimens tested to reach 20 positive specimens 21  This updated new positivity rate should be used as PIndividual in the future monitoring.

70

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Explanation: This calculation attempts to estimate the individual testing positivity rate in the  population from which the 10 (or 20) individual positive samples were collected for re assessment. Since the total number of samples (N*) that needed to be tested to obtain 10 (or

20) consecutive positive samples for re-assessment is stopped at the 10th (or 20th) positive  sample, then the positivity rate of 10/N* (or 20/N*) may be overestimated. PIndividual is corrected

in the above calculations by including an appropriate multiplier: estimated positivity rate for 10

positive sample re-assessment data is (10/N*)×(10/11) and for 20 samples is (20/N*)×(20/21).

Re-assessment of Pooling:

Re-assessment of pooling should be conducted as indicated by monitoring activities. Two possible  approaches are presented below.

Option 1

Incoming patient specimens should be tested individually until the collection contains 10 positive

specimens.

• Using these specimens, 10 n-specimen pools should be created and tested, each with 1 positive

and (n-1) negative specimens.

• Calculate the PPA between individual testing results and pooled testing results.

Note: For the calculation of the PPA, all pools that generate positive or inconclusive result are  considered in agreement with the individually-tested positive result.

Option 2

Incoming patient specimens can continue to be tested in pools.

• Re-assessment study should start from time T0 and should consist of individual sample testing in  parallel with the pooled testing. However, since all non-negative sample pools require individual  testing of all individual samples included in the pool as a part of the n-sample pooling and

deconvoluting workflow, the re-assessment study essentially consists of testing individual  samples from the negative n-sample pools.

• Re-assessment study may pause at time T1 when a minimum of 10 consecutive positive

individual results are obtained, including both positive individual results generated from  individual testing of samples from the non-negative sample pools following the n-sample pooling  and deconvoluting workflow, and positive individual results obtained from individual testing of  samples from the negative sample pools for the time period from T0 to T1 [T0, T1].

• Considering that number of positive individual sample results among negative pools is K, PPA  between testing n-sample pools and assaying single specimens using the candidate test should be

calculated as PPA (EUA Test pool vs. EUA Test individual) = 100% x (10-K)/10. It is critical that all  consecutive positive samples from time period [T0, T1] are included in the PPA calculations. With

regard to calculating the PPA, all non-negative results testing pooled samples should be counted

as in agreement with positive individually tested results.

71

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Re-assessment Acceptance Criteria for Option 1 and Option 2

• If the PPA between pooled-testing results and individual-testing results is ≥ 90% (9 out of 10 or 10

out of 10), then continuation of testing using n-specimen pooling is acceptable.  • If the PPA between pooled-testing results and individual-testing results is less than 85%:  o If PPA ≤ 70% (7 out of 10), reduce the pool size (consider n = 3) and repeat the re

assessment testing with the new pool size until PPA of pooled compared to individual  testing is ≥ 90% OR consider cessation of pooling patient specimens.

o If PPA is 80% (8 out of 10), collect an additional 10 consecutive individually positive  samples. Then, calculate the PPA from the combined data of 20 samples, between pooled

testing results and individual testing results. If the PPA is ≥ 85%, then implementation of  testing using n-sample pooling is acceptable. Or, to compensate for lost sensitivity, reduce

the pool size (consider n = 3) and repeat the re-assessment testing with the new pool size  until PPA of pooled compared to individual testing is ≥ 85%.

• If PPA of at least 85% cannot be reached for any pool size, cease pooling patient specimens.  • If n-sample pooling is acceptable, return to step 5 in “Data Monitoring” to re-establish PIndividual.

Additional Resources:

1. Abdalhamid, et al., Assessment of Specimen Pooling to Conserve SARS CoV-2 Testing resources,  Amer J Clin Pathol., Vol 153, June 2020, Pages 715–718. https://doi.org/10.1093/ajcp/aqaa064.

2. Advisory on feasibility of using pooled samples for molecular testing of COVID-19.

https://www.icmr.gov.in/pdf/covid/strategy/Advisory_on_feasibility_of_sample_pooling.pdf

3. CDC, Interim Guidance for Use of Pooling Procedures in SARS-CoV-2 Diagnostic, Screening and  Surveillance Testing. https://www.cdc.gov/coronavirus/2019-ncov/lab/pooling-procedures.html

Questions and Comments:

If you have questions or comments about this procedure, please send by email to: respvirus@cdc.gov.

72

CDC-006-00019, Revision: 06 CDC/DDID/NCIRD/ Division of Viral Diseases Effective: 12/01/2020

Division of Viral Diseases / Respiratory Viruses Branch

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel – Verification Requirements

*** DO NOT DISCARD: Important product-specific information ***

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel –  Verification Requirements

Please consult the following guidance from the Centers for Medicare & Medicaid Services  (CMS) regarding diagnostic tests under Emergency Use Authorization (EUA):

https://www.cms.gov/Medicare/Provider-Enrollment-and

Certification/SurveyCertificationGenInfo/Policy-and-Memos-to-States-and-Regions

Items/QSO18-19-CLIA

INTENDED USE

The CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel is a real time RT-PCR test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in  upper and lower respiratory specimens (such as nasopharyngeal or oropharyngeal swabs,  sputum, lower respiratory tract aspirates, bronchoalveolar lavage, and nasopharyngeal  wash/aspirate or nasal aspirate) collected from individuals suspected of COVID-19 by their  healthcare provider1.

This test is also for the qualitative detection of nucleic acid from the SARS-CoV-2 in pooled  samples containing up to four of the individual upper respiratory swab specimens

(nasopharyngeal (NP), oropharyngeal (OP), NP/OP combined, or nasal swabs) that were  collected using individual vials containing transport media from individuals suspected of  COVID-19 by their healthcare provider. Negative results from pooled testing should not be  treated as definitive. If a patient’s clinical signs and symptoms are inconsistent with a negative  result or results are necessary for patient management, then the patient should be considered  for individual testing. Specimens included in pools with a positive, inconclusive, or invalid  result must be tested individually prior to reporting a result. Specimens with low viral loads  may not be detected in sample pools due to the decreased sensitivity of pooled testing.

Testing is limited to laboratories certified under the Clinical Laboratory Improvement

Amendments of 1988 (CLIA), 42 U.S.C. § 263a, that meet the requirements to perform high  complexity tests.

Results are for the identification of SARS-CoV-2 RNA. SARS-CoV-2 RNA is generally detectable  in upper and lower respiratory specimens during infection. Positive results are indicative of  active infection with SARS-CoV-2 but do not rule out bacterial infection or co-infection with  other viruses. The agent detected may not be the definite cause of disease. Laboratories  within the United States and its territories are required to report all results to the appropriate  public health authorities.

Negative results do not preclude SARS-CoV-2 infection and should not be used as the sole  basis for treatment or other patient management decisions. Negative results must be

combined with clinical observations, patient history, and epidemiological information.

Testing with the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel is intended for use by  trained laboratory personnel who are proficient in performing real-time RT-PCR assays. The  CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel is only for use  under a Food and Drug Administration’s Emergency Use Authorization.

REQUIRED MATERIALS

The 2019 novel coronavirus positive control (nCoVPC) is provided with the CDC 2019-nCoV  Real-Time RT-PCR Diagnostic Panel and should be prepared according to the Instructions for

1 For this EUA, a healthcare provider includes, but is not limited to, physicians, nurses, pharmacists,  technologists, laboratory directors, epidemiologists, or any other practitioners or allied health

professionals.

| Document #: CDC-006-00005 | Revision #: 06 | Effective Date: 12/01/2020 | Page 1 of 5 |
| --- | --- | --- | --- |

Division of Viral Diseases / Respiratory Viruses Branch

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel – Verification Requirements

*** DO NOT DISCARD: Important product-specific information ***

Use. The nCoVPC consists of an RNA transcript of the 2019-nCoV N gene as well as human

RNase P gene segment. nCoVPC will yield a positive result with the following primer and probe  sets: 2019-nCoV_N1, 2019-nCoV_N2, and RP.

Approximately 2 mL of an upper respiratory specimen (e.g. nasopharyngeal swab (NPS) in  transport media) are needed for testing. Specimens may be pooled if less than 2 mL of one  specimen is available.

Refer to CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel package insert (manufacturer  instructions) for additional reagents, materials, and instructions.

PRECAUTIONS

This reagent should be handled in an approved biosafety level 2 (BSL-2) handling area to  avoid contamination of laboratory equipment and reagents that could cause false positive  results. This product is an RNA transcript and is non-infectious. However, the nCoVPC should  be handled in accordance with Good Laboratory Practices.

Store reagent at appropriate temperatures (see Instructions for Use) and hold on ice when  thawed.

Please use standard precautions when handling respiratory specimens.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE

QIAamp® DSP VIRAL RNA MINI KIT OR THE QIAamp® VIRAL RNA MINI KIT

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 560 ∝L of lysis buffer into each of nine tubes labeled 1-9.

• Add 140 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 14 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 14 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

• To prepare negative samples, spike 14 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE QIAGEN  EZ1® ADVANCED XL

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 280 ∝L of lysis buffer into each of nine tubes labeled 1-9.

• Add 120 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 12 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 12 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

• To prepare negative samples, spike 12 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

| Document #: CDC-006-00005 | Revision #: 06 | Effective Date: 12/01/2020 | Page 2 of 5 |
| --- | --- | --- | --- |

Division of Viral Diseases / Respiratory Viruses Branch

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel – Verification Requirements

*** DO NOT DISCARD: Important product-specific information ***

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE ROCHE  MagNA PURE TOTAL NUCLEIC ACID KIT OR THE ROCHE MagNA PURE NUCLEIC ACID

ISOLATION KIT I

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 300 ∝L of lysis buffer into each of nine tubes labeled 1-9.

• Add 100 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 12 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 12 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

• To prepare negative samples, spike 12 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE ROCHE

MagNA PURE 24 AND TOTAL NUCLEIC ACID ISOLATION KIT

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 400 ∝L of lysis buffer into each of nine tubes labeled 1-9.

• Add 100 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 12 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 12 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

• To prepare negative samples, spike 12 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE ROCHE  MagNA PURE 96 DNA AND VIRAL NA SMALL VOLUME KIT

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 350 ∝L of lysis buffer into each of nine tubes labeled 1-9.

• Add 100 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 12 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 12 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

| Document #: CDC-006-00005 | Revision #: 06 | Effective Date: 12/01/2020 | Page 3 of 5 |
| --- | --- | --- | --- |

Division of Viral Diseases / Respiratory Viruses Branch

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel – Verification Requirements

*** DO NOT DISCARD: Important product-specific information ***

• To prepare negative samples, spike 12 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE

PROMEGA MAXWELL® RSC 48

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 330 ∝L of lysis buffer (300 µL of lysis buffer + 30 µL Proteinase K, included in the  kit) into each of nine tubes labeled 1-9.

• Add 120 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 12 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 12 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

• To prepare negative samples, spike 12 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

INSTRUCTIONS FOR PREPARING SAMPLES BEFORE EXTRACTION WITH THE

BIOMÉRIEUX NucliSENS easyMAG OR THE BIOMÉRIEUX EMAG

• Refer to the 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for

reconstitution of the materials for use. RNA should be kept cold during preparation and use.  • Make a 1/10 dilution of nCoVPC by adding 5 ∝L of nCoVPC into 45 ∝L of nuclease-free water  or 10 mM Tris.

• Aliquot 1000 μL or 2000 ∝L of pre-aliquoted easyMAG lysis buffer into each of nine tubes  labeled 1-9 for the easyMAG or eMAG, respectively.

• Add 100 ∝L of upper respiratory specimen (e.g. NPS in viral transport media) into each of  the nine labeled tubes with lysis buffer.

• To prepare samples at a moderate concentration, spike 12 ∝L of undiluted nCoVPC

(rehydrated as described in the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

Instructions for Use) into each tube labeled 1-3 containing lysis buffer and specimen.

• To prepare samples at a low concentration, spike 12 ∝L of 1/10 dilution of nCoVPC into  each tube labeled 4-6 containing lysis buffer and specimen.

• To prepare negative samples, spike 12 ∝L of nuclease-free water into each tube labeled 7-9  containing lysis buffer and specimen.

• Perform extractions of all nine samples according to the CDC 2019-nCoV Real-Time RT-PCR  Diagnostic Panel Instructions for Use.

PROCEDURE

Follow the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel Instructions for Use for testing  the nine extracted samples at least once.

EXPECTED RESULTS

Moderate nCoVPC samples should be positive for 2019-nCoV.

Low nCoVPC samples should be positive for 2019-nCoV.

Negative upper respiratory samples should be negative for 2019-nCoV.

≥90% of test results should be in agreement with the expected results. If test results are  <90% in agreement with expected results, contact CDC at respvirus@cdc.gov.

| Document #: CDC-006-00005 | Revision #: 06 | Effective Date: 12/01/2020 | Page 4 of 5 |
| --- | --- | --- | --- |

Division of Viral Diseases / Respiratory Viruses Branch

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel – Verification Requirements

*** DO NOT DISCARD: Important product-specific information ***

LIMITATIONS

This test has not been FDA cleared or approved.

This test has been authorized by FDA under an EUA for use by authorized laboratories.

This test has been authorized only for the detection of nucleic acid from 2019-nCoV, not for  any other viruses or pathogens.

This test is only authorized for the duration of the declaration that circumstances exist

justifying the authorization of emergency use of in vitro diagnostics for detection and/or  diagnosis of SARS-CoV-2 under Section 564(b)(1) of the Federal Food, Drug, and Cosmetic  Act, 21 U.S.C. § 360bbb-3(b)(1), unless the authorization is terminated or revoked sooner.

QUESTIONS

Please send questions or comments by email to respvirus@cdc.gov.

DISTRIBUTION

Distributed to qualified laboratories by Centers for Disease Control and Prevention, 1600  Clifton Road, Atlanta, GA, 30329 USA

| Document #: CDC-006-00005 | Revision #: 06 | Effective Date: 12/01/2020 | Page 5 of 5 |
| --- | --- | --- | --- |

| Division of Viral Diseases/Respiratory Viruses Branch |  |
| --- | --- |
| 2019-nCoVEUA-01 | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel Product Information Sheet |

***DO NOT DISCARD: Important product-specific information***

CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel

For use under EMERGENCY USE AUTHORIZATION (EUA) only.  Rx only

CATALOG: 2019-nCoVEUA-01

KIT LOT:

EXPIRATION DATE:

INTENDED USE

The CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel is a real-time RT-PCR test intended

for the qualitative detection of nucleic acid from SARS-CoV-2 in upper and lower respiratory specimens (such as

nasopharyngeal or oropharyngeal swabs, sputum, lower respiratory tract aspirates, bronchoalveolar lavage, and

nasopharyngeal wash/aspirate or nasal aspirate) collected from individuals suspected of COVID-19 by their healthcare

provider1.

This test is also for the qualitative detection of nucleic acid from the SARS-CoV-2 in pooled samples containing up to

four of the individual upper respiratory swab specimens (nasopharyngeal (NP), oropharyngeal (OP), NP/OP combined,

or nasal swabs) that were collected using individual vials containing transport media from individuals suspected of

COVID-19 by their healthcare provider. Negative results from pooled testing should not be treated as definitive. If a

patient’s clinical signs and symptoms are inconsistent with a negative result or results are necessary for patient

management, then the patient should be considered for individual testing. Specimens included in pools with a positive,

inconclusive, or invalid result must be tested individually prior to reporting a result. Specimens with low viral loads may

not be detected in sample pools due to the decreased sensitivity of pooled testing.

Testing is limited to laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42

U.S.C. § 263a, that meet the requirements to perform high complexity tests.

Results are for the identification of SARS-CoV-2 RNA. SARS-CoV-2 RNA is generally detectable in upper and lower

respiratory specimens during infection. Positive results are indicative of active infection with SARS-CoV-2 but do not

rule out bacterial infection or co-infection with other viruses. The agent detected may not be the definite cause of

disease. Laboratories within the United States and its territories are required to report all results to the appropriate

public health authorities.

Negative results do not preclude SARS-CoV-2 infection and should not be used as the sole basis for treatment or other

patient management decisions. Negative results must be combined with clinical observations, patient history, and

epidemiological information.

Testing with the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel is intended for use by trained laboratory personnel

who are proficient in performing real-time RT-PCR assays. The CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time

RT-PCR Diagnostic Panel is only for use under a Food and Drug Administration’s Emergency Use Authorization.

PACKAGE CONTENTS

| PACKAGING | COMPONENT | PART NUMBER | COMPONENT LOT NUMBER | VIALS PER KIT | QUANTITY /VIAL | STATE |
| --- | --- | --- | --- | --- | --- | --- |
| Oligonucleotide Box | 2019-nCoV_N1 Combined Primer/Probe Mix |  |  | 1 | 22.5 nmol | Dried |
|  | 2019-nCoV_N2 Combined Primer/Probe Mix |  |  | 1 | 22.5 nmol | Dried |
|  | RP Combined Primer/Probe Mix |  |  | 1 | 22.5 nmol | Dried |
| Control Box | nCoVPC 2019-nCoV Positive Control (non-infectious) |  |  | 4 | 1 x 104 copies/µL | Dried |

IVD

1 For this EUA, a healthcare provider includes, but is not limited to, physicians, nurses, pharmacists, technologists,  laboratory directors, epidemiologists, or any other practitioners or allied health professionals.

| Document #: CDC-006-00006 | Revision #: 03 | Effective Date: 12/01/2020 | Page 1 of 2 |
| --- | --- | --- | --- |

| Division of Viral Diseases/Respiratory Viruses Branch |  |
| --- | --- |
| 2019-nCoVEUA-01 | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel Product Information Sheet |

STORAGE INSTRUCTIONS

Upon receipt, store at 2-8°C. Refer to the CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic

Panel Instructions for Use before opening and preparing reagents for use.

PROCEDURE/INTERPRETATION/LIMITATIONS

Users should refer to the CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel

Instructions for Use posted on the FDA website for all IVD products used under Emergency Use Authorization,

http://www.fda.gov/MedicalDevices/Safety/EmergencySituations/ucm161496.htm.

This test has not been FDA cleared or approved.

This test has been authorized by FDA under an EUA for use by authorized laboratories.

This test has been authorized only for the detection of nucleic acid from 2019-nCoV, not for any other viruses or

pathogens.

This test is only authorized for the duration of the declaration that circumstances exist justifying the authorization of

emergency use of in vitro diagnostics for detection and/or diagnosis of SARS-CoV-2 under Section 564(b)(1) of the

Federal Food, Drug, and Cosmetic Act, 21 U.S.C. § 360bbb-3(b)(1), unless the authorization is terminated or revoked

sooner.

PRECAUTIONS

This reagent should be handled in an approved BSL-2 handling area to avoid contamination of laboratory equipment

and reagents that could cause false positive results. This product is non-infectious. However, this product should be

handled in accordance with Good Laboratory Practices.

REAGENT COMPLAINTS/QUESTIONS

If you have a question/comment about this product, please contact the CDC Division of Viral Diseases/Respiratory

Viruses Branch by email at respvirus@cdc.gov.

DISTRIBUTED BY

Manufactured by the Centers for Disease Control and Prevention, 1600 Clifton Road, Atlanta, Georgia, 30329, USA

IVD

| Document #: CDC-006-00006 | Revision #: 03 | Effective Date: 12/01/2020 | Page 2 of 2 |
| --- | --- | --- | --- |